中国合格评定国家认可委员会

认 可 证 书 附 件

（注册号：CNAS L4802）

名称: 深圳新广行检测技术有限公司

地址：广东省深圳市南山区西丽街道西丽新围旺棠工业区18栋216

签发日期：2016年09月26日     有效期至：2022年09月25日

附件1认可的实验室关键场所一览表

| 关键场所 | 地址代码 | 地址/邮编 | 设施特点 | 主要活动 | 说明 |
| --- | --- | --- | --- | --- | --- |
|  | A | 深圳市南山区西丽新围旺棠工业区18栋216 | Ⅰ、Ⅱ | ⑴、⑵、⑶、⑷、⑸、⑹ |  |

注：

1. 设施特点包括 Ⅰ 固定、Ⅱ、离开固定设施、 Ⅲ 临时、Ⅳ 可移动、Ⅴ其他。

2. 主要活动包括 ⑴ 检测、⑵ 校准、⑶ 签发报告和/或证书、⑷ 样品接收、⑸ 合同评审、⑹ 其他。

CHINA NATIONAL ACCREDITATION SERVICE FOR CONFORMITY ASSESSMENT

SCHEDULE OF ACCREDITATION CERTIFICATE

(Registration No. CNAS L4802)

NAME: Shenzhen Xinguanghang Testing  Technology Co., Ltd.

ADDRESS: Unit 216 Room., No.18, Xili Xinwei, Redstar Community, Xili Street, Nanshan District, Shenzhen, Guangdong, China

Date of Issue:2016-09-26                      Date of Expiry:2022-09-25

SCHEDULE 1  ACCREDITED KEY LOCATIONS OF THE LABORATORY

| Locations Specified | Location Code | Address/Postal Code | Facilities Characteristic | Activity | Note |
| --- | --- | --- | --- | --- | --- |
|  | A | Unit 216 Room., No.18, Xili Xinwei, Redstar Community, Xili Street, Nanshan District, Shenzhen/518055 | Ⅰ、Ⅱ | ⑴、⑵、⑶、⑷、⑸、⑹ |  |

Note：

1. Facilities Characteristic:

Ⅰ Fixed  Facilities, Ⅱ Out of  Fixed  Facilities,  Ⅲ  Temporary Facilities,  Ⅳ  Mobile Facilities, Ⅴ Others.

2. Activity :

⑴ Testing, ⑵ Calibration, ⑶ Issue the Report and/or Certificate, ⑷ Sample Receiving, ⑸ Contract Review, ⑹ Others.

中国合格评定国家认可委员会

认 可 证 书 附 件

（注册号：CNAS L4802）

名称: 深圳新广行检测技术有限公司

地址：广东省深圳市南山区西丽街道西丽新围旺棠工业区18栋216

签发日期：2016年09月26日     有效期至：2022年09月25日

| 序号 | 姓 名 | 授权签字领域 | 备 注 |
| --- | --- | --- | --- |
| 1 | 王宁 | 全部校准、检测项目 |  |
| 2 | 杨景利 | 长度、力学、热学校准项目 |  |
| 3 | 蔡新德 | 电磁、无线电校准项目 |  |
| 4 | 周肸 | 理化、光学校准项目 |  |
| 5 | 李志伟 | 理化、光学、热工校准项目 |  |

CHINA NATIONAL ACCREDITATION SERVICE FOR CONFORMITY ASSESSMENT

SCHEDULE OF ACCREDITATION CERTIFICATE

(Registration No. CNAS L4802)

NAME: Shenzhen Xinguanghang Testing  Technology Co., Ltd.

ADDRESS: Unit 216 Room., No.18, Xili Xinwei, Redstar Community, Xili Street, Nanshan District, Shenzhen, Guangdong, China

Date of Issue:2016-09-26                      Date of Expiry:2022-09-25

SCHEDULE 2   ACCREDITED SIGNATORIES AND SCOPE

| № | Name | Authorized Scope of Signature | Note |
| --- | --- | --- | --- |
| 1 | Wang Ning | Calibration for Length,Mechanics，Thermotics Electromagnetic,Radio |  |
| 2 | Yang Jingli | Calibration for Length,Mechanics |  |
| 3 | Cai Xinde | Calibration for Electromagnetic,Radio |  |
| 4 | Zhou Xi | Calibration and Testing items including physic-chemical、photics |  |
| 5 | Li zhi wei | Calibration and Testing items including physic-chemical、photics |  |

中国合格评定国家认可委员会

认 可 证 书 附 件

（注册号：CNAS L4802）

名称: 深圳新广行检测技术有限公司

地址：广东省深圳市南山区西丽街道西丽新围旺棠工业区18栋216

认可依据：ISO/IEC 17025以及CNAS特定认可要求

签发日期：2016年09月26日     有效期至：2022年09月25日

附件3  认可的检测能力范围

| 序号 | 检测对象 | 项目/参数 |  | 领域代码 | 检测标准（方法）名称及编号（含年号） | 说明 | 备注 |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | 序号 | 名称 |  |  |  |  |
| 1 | 洁净工作台 | 1 | 洁净度 | 081202 | 层流洁净工作台检验标准 JG/T19-1999 3.2 |  |  |
|  |  | 2 | 风速 |  | 层流洁净工作台检验标准 JG/T19-1999 3.3 |  |  |
|  |  | 3 | 噪声 |  | 层流洁净工作台检验标准 JG/T19-1999 3.5 |  |  |
|  |  | 4 | 照度 |  | 层流洁净工作台检验标准 JG/T19-1999 3.7 |  |  |
| 2 | 漆膜耐冲击试验器 | 1 | 长度 | 030104 | 膜耐冲击测定法第3.2条款 冲击试验器 GB/T1732-1993 3.2 | （50±0.1）cm |  |
|  |  | 2 | 力值 |  | 膜耐冲击测定法第3.2条款 冲击试验器 GB/T1732-1993 3.2 | （1000±1）g |  |
| 3 | 岩石平板 | 1 | 平面度 | 031904 | 岩石平板 GB/T20428-2006 2.1（附录B） |  |  |
| 4 | 温度试验设备 | 1 | 温度偏差 | 031910 | 电工电子产品环境试验设备检验方法 温度试验设备 GB/T5170.2-2008 8.2.1 | 只测（-30～250）℃ | 扩项 |
|  |  | 2 | 温度波动度 |  | 电工电子产品环境试验设备检验方法 温度试验设备 GB/T5170.2-2008 8.2.1 |  | 扩项 |
|  |  | 3 | 温度均匀度 |  | 电工电子产品环境试验设备检验方法 温度试验设备 GB/T5170.2-2008 8.2.1 |  | 扩项 |
| 5 | 湿热试验设备 | 1 | 温度偏差 | 031910 | 电工电子产品环境试验设备检验方法 湿热试验设备 GB/T5170.5-2008 8.1 | 只测（-30～250）℃ | 扩项 |
|  |  | 2 | 温度波动度 |  | 电工电子产品环境试验设备检验方法 湿热试验设备 GB/T5170.5-2008 8.1 |  | 扩项 |
|  |  | 3 | 温度均匀度 |  | 电工电子产品环境试验设备检验方法 湿热试验设备 GB/T5170.5-2008 8.1 |  | 扩项 |
|  |  | 4 | 相对湿度偏差 |  | 电工电子产品环境试验设备检验方法 湿热试验设备 GB/T5170.5-2008 8.1 | 只测20%RH～98%RH | 扩项 |
|  |  | 5 | 相对湿度波动度 |  | 电工电子产品环境试验设备检验方法 湿热试验设备 GB/T5170.5-2008 8.1 |  | 扩项 |
|  |  | 6 | 相对湿度均匀度 |  | 电工电子产品环境试验设备检验方法 湿热试验设备 GB/T5170.5-2008 8.1 |  | 扩项 |
| 6 | 盐雾试验设备 | 1 | 温度偏差 | 031910 | 电工电子产品环境试验设备检验方法 盐雾试验设备 GB/T5170.8-2008 8.1 |  | 扩项 |
|  |  | 2 | 温度波动度 |  | 电工电子产品环境试验设备检验方法 盐雾试验设备 GB/T5170.8-2008 8.1 |  | 扩项 |
|  |  | 3 | 温度均匀度 |  | 电工电子产品环境试验设备检验方法 盐雾试验设备 GB/T5170.8-2008 8.1 |  | 扩项 |
|  |  | 4 | 盐雾沉降率 |  | 电工电子产品环境试验设备检验方法 盐雾试验设备 GB/T5170.8-2008 8.1 |  | 扩项 |
| 7 | 工业玻璃温度计和实验室温度计 | 1 | 温度示值 | 031908 | 工业玻璃温度计和实验室温度计 JB/T 9262-1999 6.4 | 只测（-30～250）℃ | 扩项 |
| 8 | 电子台案秤 | 1 | 称量性能 | 031912 | 电子台案秤 GB/T7722-2005 | Max≤1000kg | 扩项 |
|  |  | 2 | 重复性 | 031912 | 电子台案秤 GB/T7722-2005 GB/T7722-2005 5.4.1 |  | 扩项 |
|  |  | 3 | 偏载 |  | 电子台案秤 GB/T7722-2005 5.4.2 |  | 扩项 |
| 9 | 电子天平 | 1 | 称量性能 | 031912 | 电子天平 GB/T26497-2011 | 只测≤1000g | 扩项 |
|  |  | 2 | 重复性 | 031912 | 电子天平 GB/T26497-2011 |  | 扩项 |
|  |  | 3 | 偏载 |  | 电子天平 GB/T26497-2011 |  | 扩项 |
| 10 | 机械天平 | 1 | 计量性能 | 031912 | 机械天平 GB/T25107-2010 | 只测≤1000g d≥0.1mg | 扩项 |
| 11 | 扭力天平 | 1 | 计量性能 | 031912 | 扭力天平 GB/T25106-2010 | Max≤20g | 扩项 |
| 12 | 一般压力表 | 1 | 基本误差 | 031916 | 一般压力表 GB/T1226-2010 | （0～60）MPa | 扩项 |
|  |  | 2 | 回程误差 |  | 一般压力表 GB/T1226-2010 |  | 扩项 |
|  |  | 3 | 轻敲位移 |  | 一般压力表 GB/T1226-2010 |  | 扩项 |
| 13 | 金属洛氏硬度计 | 1 | 硬度示值 | 031919 | 金属洛氏硬度试验 第2部分 硬度计(A、B、C、D、E、F、G、H、K、N、T标尺)的检验与校准 JJG112-2013 GB/T 230.2-2002 |  | 扩项 |
|  |  | 2 | 示值重复性 |  | 金属洛氏硬度试验 第2部分 硬度计(A、B、C、D、E、F、G、H、K、N、T标尺)的检验与校准 JJG112-2013 GB/T 230.2-2003 |  | 扩项 |
| 14 | 金属维氏硬度计 | 1 | 硬度示值 | 031919 | 金属维氏硬度试验 第2部分：硬度计的检验 JJG151-2006 GB/T4340.2-1999 |  | 扩项 |
|  |  | 2 | 示值重复性 |  | 金属维氏硬度试验 第2部分：硬度计的检验 JJG151-2006 GB/T4340.2-1999 |  | 扩项 |
| 15 | 屏蔽箱（室） | 1 | 屏蔽效能 | 0430 | 电磁屏蔽室屏蔽效能的测量方法 GB/T 12190-2006 | 只测(0.5~4)GHz | 扩项 |
| 16 | LCR测试仪 | 1 | 频率 | 031924 | 军用LCR自动测量仪通用规范 SJ20842-2002 4.7.17.1.1 | 只测100Hz~10kHz | 扩项 |
|  |  | 2 | 电压幅度 |  | 军用LCR自动测量仪通用规范 SJ20842-2002 4.7.17.1.2 | 只测0.01V~5V | 扩项 |
|  |  | 3 | 电阻 |  | 军用LCR自动测量仪通用规范 SJ20842-2002 4.7.17.5.1 | 只测0.01Ω~190kΩ | 扩项 |
|  |  | 4 | 电感 |  | 军用LCR自动测量仪通用规范 SJ20842-2002 4.7.17.5.1 | 只测1μH~100μH | 扩项 |
|  |  | 5 | 电容 |  | 军用LCR自动测量仪通用规范 SJ20842-2002 4.7.17.5.1 | 只测10pF~100μF | 扩项 |
| 17 | 可见分光光度计 | 1 | 波长准确度及重复性 | 031931 | 可见分光光度计 JB/T9324-1999 |  |  |
|  |  | 2 | 透射比准确度及重复性 |  | 可见分光光度计 JB/T9324-1999 |  |  |
|  |  | 3 | 透射比与吸收比转换误差 |  | 可见分光光度计 JB/T9324-1999 |  |  |
|  |  | 4 | 杂散光 |  | 可见分光光度计 JB/T9324-1999 |  |  |

CHINA NATIONAL ACCREDITATION SERVICE FOR CONFORMITY ASSESSMENT

SCHEDULE OF ACCREDITATION CERTIFICATE

(Registration No. CNAS L4802)

Name: Shenzhen Xinguanghang Testing Technology Co., Ltd

Add：Room202-16, 2nd floor Building 18, Xinwei Wangtang Industrial Zone, Xili Street, Nanshan District, Shenzhen, Guangdong, China

Accreditation Criteria: ISO/IEC 17025 and relevant requirements of CNAS

Date of Issue:2016-09-26                      Date of Expiry:2022-09-25

SCHEDULE 3  ACCREDITED TESTING SCOPE

| №. | Test object | Item/ Parameter |  | Code of field | Title， Code of standard or method | Note | Expansion or change |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | №. | Item/ Parameter |  |  |  |  |
| 1 | Clean Benches | 1 | Cleanliness | 081202 | Test standard for laminar clean bench JG/T19-1999 3.2 |  |  |
|  |  | 2 | Wind velocity |  | Test standard for laminar clean bench JG/T19-1999 3.3 |  |  |
|  |  | 3 | Noise |  | Test standard for laminar clean bench JG/T19-1999 3.5 |  |  |
|  |  | 4 | Illuminance |  | Test standard for laminar clean bench JG/T19-1999 3.7 |  |  |
| 2 | Paint film impact resistance tester | 1 | Length | 030104 | Film shock resistance test in clause 3.2 impact tester GB/T1732-1993 3.2 |  |  |
|  |  | 2 | force |  | Film shock resistance test in clause 3.2 impact tester GB/T1732-1993 3.2 |  |  |
| 3 | granite surface plate | 1 | flatness | 031904 | granite surface plate GB/T20428-2006 2.1（附录B） |  |  |
| 4 | temperature tester | 1 | temperature deviation | 031910 | Inspection methods for environmental testing equipments for electronic and electronic products -Temperature testing equipments GB/T5170.2-2008 8.2.1 | Accredited only for（-30～250）℃ | Expansion |
|  |  | 2 | temperature fluctuation |  | Inspection methods for environmental testing equipments for electronic and electronic products -Temperature testing equipments GB/T5170.2-2008 8.2.1 |  | Expansion |
|  |  | 3 | temperature uniformity |  | Inspection methods for environmental testing equipments for electronic and electronic products -Temperature testing equipments equipment testing equipmen GB/T5170.2-2008 8.2.1 |  | Expansion |
| 5 | Damp heat testing equipment | 1 | temperature deviation | 031910 | Inspection methods for environmental testing equipments for electronic and electronic products-Damp heat testing equipments GB/T5170.5-2008 8.1 | Accredited only for（-30～250）℃ | Expansion |
|  |  | 2 | temperature fluctuation |  | Inspection methods for environmental testing equipments for electronic and electronic products-Damp heat testing equipments GB/T5170.5-2008 8.1 |  | Expansion |
|  |  | 3 | temperature uniformity |  | Inspection methods for environmental testing equipments for electronic and electronic products-Damp heat testing equipments GB/T5170.5-2008 8.1 |  | Expansion |
|  |  | 4 | Relative humidity deviation |  | Inspection methods for environmental testing equipments for electronic and electronic products-Damp heat testing equipments GB/T5170.5-2008 8.1 | Accredited only for 20%RH～98%RH | Expansion |
|  |  | 5 | Relative humidity fluctuation |  | Inspection methods for environmental testing equipments for electronic and electronic products-Damp heat testing equipments GB/T5170.5-2008 8.1 |  | Expansion |
|  |  | 6 | Relative humidity uniformity |  | Inspection methods for environmental testing equipments for electronic and electronic products-Damp heat testing equipments GB/T5170.5-2008 8.1 |  | Expansion |
| 6 | Salt spray testing equipment | 1 | temperature deviation | 031910 | Inspection methods for environmental testing equipments for electronic and electronic products-Salt mist testing equipments GB/T5170.8-2008 8.1 |  | Expansion |
|  |  | 2 | temperature fluctuation |  | Inspection methods for environmental testing equipments for electronic and electronic products-Salt mist testing equipments GB/T5170.8-2008 8.1 |  | Expansion |
|  |  | 3 | temperature uniformity |  | Inspection methods for environmental testing equipments for electronic and electronic products-Salt mist testing equipments GB/T5170.8-2008 8.1 |  | Expansion |
|  |  | 4 | Salt fog sedimentation rate |  | Inspection methods for environmental testing equipments for electronic and electronic products-Salt mist testing equipments GB/T5170.8-2008 8.1 |  | Expansion |
| 7 | Industrial glass thermometer and laboratory thermometer | 1 | Temperature value | 031908 | Liquid-in-glass industrial thermometers and laboratorial thermometers JB/T 9262-1999 6.4 | Accredited only for（-30～250）℃ | Expansion |
| 8 | Electronic portable plateform and bench scale | 1 | Weighing performance | 031912 | Electronic portable plateform and bench scale GB/T7722-2005 | Max≤1000kg | Expansion |
|  |  | 2 | Repeatability | 031912 | Electronic portable plateform and bench scale GB/T7722-2005 5.4.1 |  | Expansion |
|  |  | 3 | Unbalance loading |  | Electronic portable plateform and bench scale GB/T7722-2005 5.4.2 |  | Expansion |
| 9 | Electronic balance | 1 | Weighing performance | 031912 | Electronic balance GB/T26497-2011 | only for≤1000g | Expansion |
|  |  | 2 |  | 031912 | Electronic balance GB/T26497-2011 |  | Expansion |
|  |  | 3 | Partial load |  | Electronic balance GB/T26497-2011 |  | Expansion |
| 10 | Mechanical balance | 1 | Measuring performance | 031912 | Mechanical balance GB/T25107-2010 | only for≤1000g d≥0.1mg | Expansion |
| 11 | Torsion balance | 1 | Measuring performance | 031912 | Torsion balance GB/T25106-2010 | Max≤20g | Expansion |
| 12 | General pressure gauge | 1 | basic error | 031916 | General pressure gauge GB/T1226-2010 | （0～60）MPa | Expansion |
|  |  | 2 | hysterisis error |  | General pressure gauge GB/T1226-2010 |  | Expansion |
|  |  | 3 | friction error |  | General pressure gauge GB/T1226-2010 |  | Expansion |
| 13 | Metallic Rockwell Hardness testing Machines | 1 | Hardness value | 031919 | Metal rockwell hardness test Part 2 hardness tester (A, B, C, D, E, F, G, H, N, K, T scale) of the test and calibration JJG112-2013 GB/T 230.2-2002 |  | Expansion |
|  |  | 2 | Indication repeatability |  | Metal rockwell hardness test Part 2 hardness tester (A, B, C, D, E, F, G, H, N, K, T scale) of the test and calibration JJG112-2013 GB/T 230.2-2003 |  | Expansion |
| 14 | Metal vickers hardness tester | 1 | Hardness value | 031919 | Metal vickers hardness test Part 2: Verification and calibration of testing machines JJG151-2006 GB/T4340.2-1999 |  | Expansion |
|  |  | 2 | Indication repeatability |  | Metal vickers hardness test Part 2: Verification and calibration of testing machines JJG151-2006 GB/T4340.2-1999 |  | Expansion |
| 15 | Shielding box (room) | 1 | Shielding effectiveness | 0430 | Measurement of shielding effectiveness of electromagnetic shielding room GB/T 12190-2006 | accredited only for(0.5~4)GHz |  |
| 16 | LCR Tester | 1 | Frequency | 031924 | Military general specification for LCR automatic measuring instrument SJ20842-2002 4.7.17.1.1 | Accredited only for100Hz~10kHz |  |
|  |  | 2 | Amplitude |  | Military general specification for LCR automatic measuring instrument SJ20842-2002 4.7.17.1.2 | Accredited only for0.01V~5V |  |
|  |  | 3 | Resistance |  | Military general specification for LCR automatic measuring instrument SJ20842-2002 4.7.17.5.1 | Accredited only for0.01Ω~190kΩ |  |
|  |  | 4 | Inductance |  | Military general specification for LCR automatic measuring instrument SJ20842-2002 4.7.17.5.1 | Accredited only for1μH~100μH |  |
|  |  | 5 | Capacitance |  | Military general specification for LCR automatic measuring instrument SJ20842-2002 4.7.17.5.1 | Accredited only for10pF~100μF |  |
| 17 | Visible spectrophotometer | 1 | Wavelength accuracy and repeatability | 031931 | Visible spectrophotometer JB/T9324-1999 |  |  |
|  |  | 2 | Transmittance accuracy and repeatability |  | Visible spectrophotometer JB/T9324-1999 |  |  |
|  |  | 3 | Transmittance and absorption than conversion error |  | Visible spectrophotometer JB/T9324-1999 |  |  |
|  |  | 4 | Stray light |  | Visible spectrophotometer JB/T9324-1999 |  |  |

中国合格评定国家认可委员会

认 可 证 书 附 件

（注册号：CNAS L4802）

名称: 深圳新广行检测技术有限公司

地址：广东省深圳市南山区西丽街道旺棠工业区18栋216

认可依据：ISO/IEC 17025以及CNAS特定认可要求

签发日期：2016年09月26日     有效期至：2022年09月25日

附件5  认可的校准和测量能力范围

注：“测量仪器名称”栏仪器名称前标注*的项目可开展现场校准。

| 序号 | 测量仪器名称 | 校准参量 | 领域代码 | 校准规范名称及编号（含年号） | 测量范围 | 扩展不确定度 (k=2) | 说明 | 备注 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 未分组 |  |  |  |  |  |  |  |  |
| 1 | 通用卡尺 | 长度 | 600213 | 通用卡尺检定规程 JJG30 | (0～300)mm (游标类) | U=0.02mm |  |  |
|  |  |  |  |  | (>300～600)mm (游标类) | U=0.04mm |  |  |
|  |  |  |  |  | (0～300)mm (带表类) | U=0.02mm |  |  |
|  |  |  |  |  | (>300～600)mm (带表类) | U=0.04mm |  |  |
|  |  |  |  |  | (0～300)mm (数显类) | U=0.01mm |  |  |
|  |  |  |  |  | (>300～600)mm (数显类) | U=0.04mm |  |  |
|  |  |  |  |  | (0～300)mm (深度类) | U=0.01mm |  |  |
| 2 | 高度卡尺 | 长度 | 600215 | 高度卡尺检定规程 JJG31 | (0～300)mm (游标类) | U=0.02mm |  |  |
|  |  |  |  |  | (>300～600)mm (游标类) | U=0.04mm |  |  |
|  |  |  |  |  | (0～300)mm (带表类) | U=0.02mm |  |  |
|  |  |  |  |  | (>300～600)mm(带表类) | U=0.04mm |  |  |
|  |  |  |  |  | (0～300)mm (数显类) | U=0.01mm |  |  |
|  |  |  |  |  | (>300～600)mm (数显类) | U=0.04mm |  |  |
| 3 | 指示表 | 长度 | 600208 | 指示表（指针式、数显式）检定规程 JJG34 | 百分表(0～3)mm | U=3.0μm |  |  |
|  |  |  |  |  | (>3～5)mm (百分表) | U=3.5μm |  |  |
|  |  |  |  |  | (>5～10)mm (百分表) | U=4.4μm |  |  |
|  |  |  |  |  | (0～1)mm (千分表) | U=1.2μm |  |  |
|  |  |  |  |  | (>1～2)mm (千分表) | U=1.8μm |  |  |
| 4 | 深度百分表 | 长度 | 600212 | 深度指示表检定规程 JJG830 | (0～10)mm | U=2.5μm |  |  |
|  |  |  |  |  | (>10～50)mm | U=2.8μm |  |  |
|  |  |  |  |  | (>50～100)mm | U=3.5μm |  |  |
| 5 | 杠杆表 | 长度 | 600209 | 杠杆表检定规程 JJG35 | 杠杆百分表： (0～0.8)mm | U=3.0μm |  |  |
|  |  |  |  |  | 杠杆千分表： (0～0.2)mm | U=1.8μm |  |  |
| 6 | 内径表 | 长度 | 600211 | 内径表校准规范 JJF1102 | 内径百分表： (6～450)mm | U=3.0μm |  |  |
|  |  |  |  |  | 内径千分表： (10～450)mm | U=1.2μm |  |  |
| 7 | 千分尺 | 长度 | 600204 | 千分尺检定规程 JJG21 | 外径类(0～50)mm | U=1.4μm |  |  |
|  |  |  |  |  | 外径类(>50～100)mm | U=1.6μm |  |  |
|  |  |  |  |  | 外径类(>100～300)mm | U=1.9μm |  |  |
|  |  |  |  |  | 数显外径类(0～50)mm | U=1.0μm |  |  |
|  |  |  |  |  | 数显外径类(>50～100)mm | U=1.2μm |  |  |
|  |  |  |  |  | 数显外径类(>100～300)mm | U=1.4μm |  |  |
| 8 | 内测千分尺 | 长度 | 600202 | 测量内尺寸千分尺校准规范 JJF1091 | (5～75)mm | U=1.6μm |  |  |
|  |  |  |  |  | (>75～125)mm | U=2.2μm |  |  |
|  |  |  |  |  | (>125～150)mm | U=2.5μm |  |  |
| 9 | 深度千分尺 | 长度 | 600205 | 深度千分尺检定规程 JJG24 | (0～50)mm | U=1.1μm |  |  |
|  |  |  |  |  | (>50～125)mm | U=1.3μm |  |  |
|  |  |  |  |  | (>125～150)mm | U=1.5μm |  |  |
|  |  |  |  |  | (>150～250)mm | U=1.7μm |  |  |
|  |  |  |  |  | (>250～300)mm | U=1.9μm |  |  |
| 10 | 公法线千分尺 | 长度 | 600511 | 公法线千分尺检定规程 JJG82 | (0～25)mm | U=0.8μm |  |  |
|  |  |  |  |  | (>25～50)mm | U=1.6μm |  |  |
|  |  |  |  |  | (>50～75)mm | U=1.9μm |  |  |
|  |  |  |  |  | (>75～100)mm | U=2.4μm |  |  |
| 11 | 量块 | 长度 | 600201 | 量块检定规程 JJG146 | (0.5～100)mm | U=(0.15+1.5L)μm(L-m) |  |  |
|  |  |  |  |  | (>100～291.8)mm | U=(0.37+3.7L)μm (L-m) |  |  |
| 12 | 投影仪 | 长度 | 600122 | 投影仪校准规范 JJF1093 | X、Y轴：0mm＜L≤200mm | U=2.0μm |  |  |
| 13 | 影像测量仪 | 长度 | 600633 | 影像测量仪校准规范 JJF1318 | X、Y轴：0mm＜L≤300mm | U=2.0μm |  |  |
| 14 | 测量显微镜 | 长度 | 600120 | 读数、测量显微镜检定规程 JJG571 | 纵向：(0～50)mm 横向：(0～13)mm | U=4.0μm |  |  |
| 15 | 磁性、电涡流式覆层厚度测量仪 | 厚度 | 600635 | 磁性、电涡流式覆层厚度测量仪检定规程 JJG818 | (0～1000)μm | U=(2+2H%)μm |  |  |
| 16 | 坐标测量机 | 长度 | 600601 | 坐标测量机校准规范 JJF1064 | X、Y、Z轴：100mm≤L≤1000mm | U=(0.6+3L／1000)μm (L-mm) |  |  |
| 17 | 工具显微镜 | 长度 | 600119 | 工具显微镜检定规程 JJG56 | X、Y-轴：0mm＜L≤200mm | U=1.3μm |  |  |
| 18 | 光滑极限量规 | 长度 | 600613 | 光滑极限量规检定规程 JJG343 | 光面塞规： Φ(1～100)mm | U=(0.5+5L)μm (L-m) |  |  |
|  |  |  |  |  | 光面环规： Φ(16～100)mm | U=(0.7+6L)μm(L-m) |  |  |
| 19 | 塞尺 | 长度 | 600608 | 塞尺检定规程 JJG62 | (0.02～0.10)mm | U=1.5μm |  |  |
|  |  |  |  |  | (>0.10～1.0)mm | U=2.0μm |  |  |
| 20 | 三针 | 长度 | 600616 | 针规、三针校准规范 JJF1207 | Φ(0.118～6.585)mm | U=0.4μm |  |  |
| 21 | 针规 | 长度 | 600616 | 针规、三针校准规范 JJF1207 | Φ(0.3～10)mm | U=1.0μm |  |  |
|  |  |  |  |  | Φ(>10～25)mm | U=1.3μm |  |  |
| 22 | 圆柱螺纹量规 | 长度 | 600612 | 圆柱螺纹量规校准规范 JJF1345 | 螺纹塞规： M1.2～M20 | U=3.6μm |  |  |
|  |  |  |  |  | 螺纹塞规： >M20～M50 | U=4.0μm |  |  |
|  |  |  |  |  | 螺纹塞规： >M50～M100 | U=4.5μm |  |  |
|  |  |  |  |  | 螺纹环规： M2～M20 | U=4.2μm |  |  |
|  |  |  |  |  | 螺纹环规： >M20～M50 | U=4.6μm |  |  |
|  |  |  |  |  | 螺纹环规： >M50～M100 | U=5.0μm |  |  |
| 23 | 橡胶测厚计 | 长度 | 600219 | 厚度表校准规范 JJF1255 | (0～10)mm | U=3μm |  |  |
| 24 | 数显测高仪 | 长度 | 600410 | 数显测高仪校准规范 JJF1245 | (0～400)mm | U=1.5μm |  |  |
|  |  |  |  |  | (>400～600)mm | U=4.0μm |  |  |
| 25 | 百分表式卡规 | 长度 | 600611 | 百分表式卡规检定规程 JJG109 | (0～50)mm | U=3.2μm |  |  |
|  |  |  |  |  | (>50～100)mm | U=3.5μm |  |  |
| 26 | 测长仪 | 长度 | 600618 | 测长仪校准规范 JJF1189 | 绝对测量： (0～100)mm | U=(0.4+L/300)μm(L-m) |  |  |
| 27 | 指示类量具检定仪 | 长度 | 600220 | 指示类量具检定仪检定规程 JJG201 | 百分表检定仪： (0～25)mm | U=0.9μm |  |  |
|  |  |  |  |  | 千分表检定仪： (0～2)mm | U=0.5μm |  |  |
| 28 | 杠杆式千分尺、杠杆式卡规 | 长度 | 600611 | 杠杆千分尺、杠杆卡规检定规程 JJG26 | (0～25)mm | U=0.9μm |  |  |
|  |  |  |  |  | (>25～50)mm | U=1.1μm |  |  |
|  |  |  |  |  | (>50～75)mm | U=1.5μm |  |  |
|  |  |  |  |  | (>75～100)mm | U=1.8μm |  |  |
| 29 | 大量程百分表 | 长度 | 600210 | 大量程百分表检定规程 JJG379 | (0～10)mm | U=4.4μm |  |  |
|  |  |  |  |  | (>10～30)mm | U=5.1μm |  |  |
|  |  |  |  |  | (>30～50)mm | U=5.8μm |  |  |
| 30 | 钢卷尺 | 长度 | 600109 | 钢卷尺检定规程 JJG4 | (0～1)m | U=0.2mm |  |  |
|  |  |  |  |  | (>1～3)m | U=0.3mm |  |  |
|  |  |  |  |  | (>3～5)m | U=0.5mm |  |  |
| 31 | 织物厚度仪 | 长度 | 700120 | 织物厚度仪校准规范 JJF (纺织)020 | (0～10)mm | U=3μm |  | 扩项 |
|  |  | 长度 |  |  | (0～300)mm | U=0.03mm |  |  |
|  |  | 力值 |  |  | (0～10)N | U=0.02N |  |  |
| 32 | 刀口形直尺 | 长度 | 600308 | 刀口形直尺检定规程 JJG63 | (75～300)mm | U=0.7μm |  | 扩项 |
| 33 | 试验筛 | 长度 | 600641 | 试验筛校准规范 JJF1175 | (2～50)mm | U=0.03mm |  | 扩项 |
| 34 | 试模 | 长度 | 600647 | 试模校准规范 JJF1307 | (300×300)mm | U=0.03 mm |  | 扩项 |
| 35 | 光栅测微仪 | 长度 | 600623 | 光栅式测微仪检定规程 JJG 989 | (0～10) mm | U=0.29 μm |  | 扩项 |
|  |  |  |  |  | (＞10～30) mm | U=0.32μm |  |  |
|  |  |  |  |  | (＞30～50) mm | U=0.33 μm |  |  |
|  |  |  |  |  | (＞50～100) mm | U=0.39 μm |  |  |
| 36 | 跳动检查仪(偏摆检查仪) | 长度 | 600508 | 跳动检查仪校准规范 JJF1109 | (1～500 )mm | U=1.5μm |  | 扩项 |
| 37 | 超声波测厚仪 | 长度 | 600635 | 超声波测厚仪校准规范 JJF1126 | (0.5～75)mm | U=32μm |  | 扩项 |
|  |  |  |  |  | (＞75～200)mm | U= 44μm |  |  |
| 38 | 电感测微仪 | 长度 | 600623 | 电感测微仪校准规范 JJF1331 | (0～1000)μm | U=0.6μm |  | 扩项 |
| 39 | 扭簧比较仪 | 长度 | 600622 | 扭簧比较仪检定规程 JJG 118 | (0～100)μm | U=0.2μm |  | 扩项 |
| 40 | 机械式比较仪 | 长度 | 600622 | 机械式比较仪检定规程 JJG39 | (0～100)μm | U=0.2μm |  | 扩项 |
| 41 | 内径千分尺 | 长度 | 600202 | 内径千分尺检定规程 JJG 22 | (50～500)mm | U=2.0μm |  | 扩项 |
| 42 | 激光测径仪 | 长度 | 600620 | 激光测径仪校准规范 JJF1250 | Φ(0.1～30)mm | U=1.3μm |  | 扩项 |
| 43 | 厚度表 | 长度 | 600219 | 厚度表校准规范 JJF1255 | (0.1～30)mm | U=0.02mm（分度值0.1mm） |  | 扩项 |
|  |  |  |  |  | (0.1～30)mm | U=2μm（分度值0.001mm） |  |  |
| 44 | X射线荧光镀层测厚仪 | 长度 | 600219 | X射线荧光镀层测厚仪校准规范 JJF1306 | Au/Ni,Sn/Cu,Cu/Fe,Ni/Fe,Ni/Cu(0～15)μm | U=(0.01+9.6H%)μm |  | 扩项 |
| 45 | 锡膏厚度测量仪 | 长度 | 600219 | X射线荧光镀层测厚仪校准规范 JJF 1306 | (40～400)μm | U=(0.01+9.6H%)μm |  | 扩项 |
| 46 | 计米器 | 长度 | 600117 | 线缆计米器检定规程 JJG 987 | (0.01～5)m | Urel=0.3% |  | 扩项 |
| 47 | 刮板细度计 | 长度 | 600629 | 刮板细度计检定规程 JJG905 | (0～50)μm | U=0.4μm |  | 扩项 |
|  |  |  |  |  | (50～150)μm | U=1.2μm |  |  |
| 48 | 万能角度尺 | 角度 | 600312 | 万能角度尺检定规程 JJG33 | (0～360)° | U=1.3′ |  | 扩项 |
| 49 | 接触式表面粗糙度测量仪 | 粗糙度 | 600626 | 触针式表面粗糙度测量仪校准规范 JJF1105 | Ra(0.2～3.2)μm | Urel=5.5% |  | 扩项 |
| 50 | *圆度、圆柱度测量仪 | 圆度 | 600602 | 圆度 圆柱度测量仪检定规程 JJG429 | (0～15)μm | U=0.2μm | 只做径向、轴向误差 | 扩项 |
| 51 | 数字温度指示（调节）仪 | 温度 | 610208 | 数字温度指示调节仪检定规程 JJG617 | RTD：(-50～0)℃ | U=(0.3～0.2)℃ |  |  |
|  |  |  |  |  | RTD：(0～800)℃ | U=(0.2～0.5)℃ |  |  |
|  |  |  |  |  | TC：(-50～0)℃ | U=(0.3～0.2)℃ |  |  |
|  |  |  |  |  | TC：(0～1300)℃ | U=(0.2～0.8)℃ |  |  |
| 52 | 模拟温度指示调节仪 | 温度 | 610207 | 模拟式温度指示调节仪检定规程 JJG 951 | RTD：(-100～0)℃ | U=(0.5～0.3)℃ |  | 扩项 |
|  |  |  |  |  | RTD：(0～800)℃ | U=(0.3～1.0)℃ |  |  |
|  |  |  |  |  | TC：(-100～0)℃ | U=(0.5～0.3)℃ |  |  |
|  |  |  |  |  | TC：(0～1300)℃ | U=(0.3～1.2)℃ |  |  |
| 53 | 温度指示控制仪 | 温度 | 610202 | 温度指示控制仪检定规程 JJG 874 | (-30～300)℃ | U=(0.3～0.8)℃ |  | 扩项 |
| 54 | 双金属温度计 | 温度 | 610116 | 双金属温度计检定规程 JJG 226 | (-30～300)℃ | U=0.3℃ |  | 扩项 |
| 55 | 温度巡回检测仪 | 温度 | 610204 | 温度巡回检测仪校准规范 JJF 1171 | (-30～300)℃ | U=(0.2～0.4)℃ |  | 扩项 |
| 56 | 工作用铜-铜镍热电偶 | 温度 | 610106 | 工作用铜-铜镍热电偶检定规程 JJG368 | (-30～300)℃ | U=(0.2～0.6)℃ |  | 扩项 |
| 57 | 工作用廉金属热电偶 | 温度 | 610106 | 工作用廉金属热电偶检定规程 JJG 351 | (-30～300)℃ | U=(0.2～0.6)℃ |  | 扩项 |
| 58 | 工业铂热电阻 | 温度 | 610110 | 工业铂、铜热电阻检定规程 JJG 229 | (-30～300)℃ | U=0.3℃ |  | 扩项 |
| 59 | 工业铜热电阻 | 温度 | 610111 | 工业铂、铜热电阻检定规程 JJG 229 | (-30～300)℃ | U=0.3℃ |  | 扩项 |
| 60 | 温湿度计 | 温度 | 610603 | 机械式温湿度计检定规程 JJG 205 | (5～50)℃ | U=0.5℃ |  | 扩项 |
|  |  | 湿度 |  |  | (20～90)%RH | U=2.0%RH |  |  |
| 61 | *恒温槽 | 温度 | 610302 | 恒温槽技术性能测试规范 JJF1030 | (-100～300)℃ | U=(0.09～0.15)℃ |  | 扩项 |
| 62 | *环境试验设备 | 温度 | 610306 | 环境试验设备温度、湿度校准规范 JJF 1101 | (室温～300)℃ | U=(0.5～1.0)℃ |  | 扩项 |
| 63 | 表面温度计 | 温度 | 610118 | 表面温度计校准规范 JJF 1409 | (室温～300) ℃ | U=(0.3～0.8)℃ |  | 扩项 |
| 64 | 工业温测量记录仪 | 温度 | 610203 | 工业过程测量记录仪检定规程 JJG 74 | RTD：(-200～+850)℃ | U=(0.6～0.9)℃ |  | 扩项 |
|  |  |  |  |  | TC：(0～1600)℃ | U=(1.0～1.2)℃ |  |  |
| 65 | 温度校准仪 | 温度 | 610201 | 温度校准仪校准规范 JJF1309 | TC测量：(-270～1372)℃ | U=(0.2～0.5)℃ |  | 扩项 |
|  |  |  |  |  | RTD测量：(-200～850)℃ | U=(0.2～0.4)℃ |  |  |
|  |  |  |  |  | TC输出：(-270～1300)℃ | U=(0.2～0.5)℃ |  |  |
|  |  |  |  |  | RTD输出：(-200～850)℃ | U=(0.2～0.4)℃ |  |  |
| 66 | 压力式温度计 | 温度 | 610117 | 压力式温度计检定规程 JJG 310 | (-30～300)℃ | U=0.5℃ |  | 扩项 |
| 67 | *盐雾试验设备 | 温度 | 610311 | 盐雾试验设备技术性能校准规范 JJF（闽）1044 | (30～95)℃ | U=0.4℃ |  | 扩项 |
|  |  | 沉降量 |  |  | (0～5)mL/(h.80cm2) | U=0.06mL/(h.80cm2) |  |  |
| 68 | *箱式电阻炉 | 温度 | 610308 | 箱式电阻炉校准规范 JJF 1376 | (室温～1300)℃ | U=(0.7～1.4)℃ |  | 扩项 |
| 69 | 热敏电阻测温仪 | 温度 | 610122 | 热敏电阻测温仪校准规范 JJF 1379 | (－30～300)℃ | U=0.3℃ |  | 扩项 |
| 70 | *热变形、维卡软化点温度测定仪 | 温度 | 610399 | 热变形、维卡软化点温度测定仪校准规范1 JJF (浙) 1051 | (室温～300)℃ | U=(0.4～0.5)℃ |  | 扩项 |
|  |  | 升温速率 | 610399 |  | (0.1～140)℃/h | U=(0.3～0.4)℃/h |  |  |
|  |  | 位移 | 610399 |  | (0.5～10)mm | U=4um |  |  |
|  |  | 力值 | 610399 |  | (0.01～50)N | U=0.02N |  |  |
| 71 | *医用热力灭菌设备温度计 | 温度 | 610199 | 医用热力灭菌设备温度计校准规范 JJF 1308 | (室温～150)℃ | U=1.0℃ |  | 扩项 |
| 72 | *氧弹、空气弹老化试验箱 | 温度 | 610314 | 橡皮塑料电线电缆试验仪器设备检定方法 第9部分：氧弹、空气弹老化试验箱 JB/T 4278.9 | (50～300)℃ | U=1.0℃ |  | 扩项 |
| 73 | 温度变送器(带传感器) | 温度 | 610210 | 温度变送器校准规范 JJF 1183 | 配热电阻(－30～250)℃ | U=0.08℃ |  | 扩项 |
|  |  |  |  |  | 配热电偶(－30～250)℃ | U=0.4℃ |  |  |
| 74 | 温度变送器(不带传感器) | 温度 | 610210 | 温度变送器校准规范 JJF 1183 | 配热电阻(－30～250)℃ | U=0.10℃ |  | 扩项 |
|  |  |  |  |  | 配热电偶(－30～250)℃ | U=0.4℃ |  |  |
| 75 | *灼热丝试验仪 | 温度 | 701049 | 灼热丝试验仪校准规范 JJF(浙)1050 | (50～1000)℃ | U=1.0℃ |  | 扩项 |
|  |  | 长度 |  |  | (0～200) mm | U=0.02mm |  |  |
|  |  | 力值 |  |  | (0～10)N | U=0.1N |  |  |
|  |  | 时间 |  |  | (1～600)s | U=0.12s |  |  |
| 76 | WBGT指数仪温度计 | 温度 | 610199 | WBGT 指数仪温度计校准规范 JJF 1407 | (5～120)℃ | U=0.3℃ |  | 扩项 |
| 77 | 熔点测定仪 | 熔点 | 681006 | 熔点测定仪检定规程 JJG 701 | (50～250)℃ | U=0.4℃ |  | 扩项 |
| 78 | *工作测力仪 | 力值 | 620799 | 工作测力仪检定规程 JJG455 | (0.03～1000)N | Urel=0.4% |  |  |
| 79 | *拉力、压力和万能试验机 | 力值 | 620710 | 拉力、压力和万能试验机检定规程 JJG139 | (0.1～200)kN | Urel=0.4% |  | 变更 |
| 80 | *转速表 | 转速 | 621102 | 转速表检定规程 JJG105 | (30～30000)r/min | Urel=0.3% |  | 扩项 |
| 81 | *扭矩扳子 | 扭矩 | 620805 | 扭矩扳子检定规程 JJG 707 | (0.1～100)Nm | Urel=1.5% |  | 扩项 |
| 82 | *扭矩仪 | 扭矩 | 620803 | 扭矩扳子检定仪检定规程 JJG 797 | (0.2～50)Nm | Urel=0.7% |  | 扩项 |
| 83 | *电子式万能试验机 | 力值 | 620711 | 电子式万能试验机检定规程 JJG 475 | 1N～200kN | Urel=0.4% |  | 扩项 |
| 84 | *弹簧拉压试验机 | 力值 | 620710 | 专用工作测力机校准规范 JJF 1134 | 1N～200kN | Urel=0.4% |  | 扩项 |
| 85 | *非金属万能试验机 | 力值 | 620711 | 拉力、压力和万能试验机检定规程 JJG 139 | 50N～200kN | Urel=0.4% |  | 扩项 |
| 86 | *水泥抗折试验机 | 力值 | 620713 | 抗折试验机检定规程 JJG 476 | (0.6～6)kN | Urel=0.4% |  | 扩项 |
| 87 | 橡胶硬度计(A型邵氏硬度计) | 力值 | 620916 | A型邵氏硬度计检定规程 JJG 304 | (1.050～8.900)N | U=29mN |  | 扩项 |
|  |  | 压针尺寸 |  |  | (0～2.5)mm | U=0.01mm |  |  |
| 88 | 橡胶硬度计(D型邵氏硬度计) | 力值 | 620916 | D型邵氏硬度计检定规程 JJG 1039 | (4～45)N | U=150mN |  | 扩项 |
|  |  | 压针尺寸 |  |  | (0～2.5)mm | U=0.01mm |  |  |
| 89 | *量仪测力仪 | 力值 | 620708 | 专用工作测力机校准规范 JJF 1134 | (0.03～15)N | Urel =0.2% |  | 扩项 |
| 90 | *纸与纸板破裂强度试验机 | 压力 | 700416 | 纸与纸板耐破度测定仪检定规程 JJG (轻工) 61 | (0.1～2000)kPa | U=0.5% |  | 扩项 |
|  |  | 长度 |  |  | (0～50)mm | U=0.10mm |  |  |
| 91 | *纸箱抗压试验机 | 力值 | 700428 | 纸箱抗压试验机 JJG (轻工) 115 | (0.1～200)kN | Urel=0.4% |  | 扩项 |
| 92 | *测功机 | 转速 | 620809 | 测功装置检定规程 JJG 653 | (5～4000)r/min | Urel=0.2% | 限于非接触式 | 扩项 |
|  |  | 扭矩 |  |  | (0.01～50)Nm | Urel=0.4% |  |  |
| 93 | *纸板压缩强度试验仪 | 力值 | 700401 | 纸板压缩强度试验仪检定规程 JJG（轻工）49 | (50～3000)N | Urel=0.4% |  | 扩项 |
| 94 | *水泥胶砂搅拌机 | 转速 | 700816 | 水泥胶砂搅拌机检定规程 JJG(建材)102 | 65r/min,137 r/min | U=1r/min |  | 扩项 |
|  |  | 时间 |  |  | (0～999)s | U =1s |  |  |
|  |  | 长度 |  |  | (0～300)mm | U = 0.03mm |  |  |
| 95 | *行星式胶砂搅拌机 | 转速 | 700816 | 行星式胶砂搅拌机检定规程 JJG(建材)123 | 140 r/min,285 r/min | U=1r/min |  | 扩项 |
|  |  | 时间 |  |  | (0～999)s | U =0.3s |  |  |
|  |  | 长度 |  |  | (0～300)mm | U = 0.3mm |  |  |
| 96 | *非金属薄板抗折试验机 | 力值 | 700820 | 抗折试验机检定规程 JJG 476 | (0.1～20)kN | Urel=0.4% |  | 扩项 |
| 97 | *水泥净浆搅拌机 | 转速 | 700815 | 水泥净浆搅拌机检定规程 JJG(建材) 104 | (62～285) r/min | U=1r/min |  | 扩项 |
|  |  | 长度 |  |  | (0～200)mm | U = 0.3mm |  |  |
|  |  | 时间 |  |  | 1s～29min | U=0.3s |  |  |
| 98 | *雷氏夹膨胀测定仪 | 长度 | 700847 | 雷氏夹膨胀测定仪检定规程 JJG(建材) 110 | (0～200)mm | U = 0.3mm |  | 扩项 |
|  |  | 质量 |  |  | 300g | U=0.03g |  |  |
| 99 | *胶砂试体成型振实台 | 振幅 | 700828 | 胶砂试体成型振实台检定规程 JJG(建材) 124 | 14.7mm～15.3mm | U = 0.1mm |  | 扩项 |
|  |  | 时间 |  |  | 60s | U=0.2s |  |  |
|  |  | 质量 |  |  | (13.5～30)kg | U=0.1kg |  |  |
| 100 | *水泥胶砂流动度测定仪 | 质量 | 700808 | 水泥胶砂流动度测定仪检定规程 JJG(建材) 126 | 3.45kg | U=2g |  | 扩项 |
|  |  | 时间 |  |  | 30s | U=0.3s |  |  |
|  |  | 长度 |  |  | 200m | U=0.1mm |  |  |
| 101 | *水泥胶砂振动台 | 转速 | 700818 | 水泥胶砂振动台检定规程 JJG 918 | (2800～3000)r/min,(46.7～50)Hz | U=1r/min,U=0.05Hz |  | 扩项 |
|  |  | 位移 |  |  | (0～50)mm | Urel=3% |  |  |
|  |  | 时间 |  |  | 1s～29min | U=0.12s |  |  |
| 102 | *耐洗色牢度试验机 | 转速 | 700127 | 耐洗色牢度试验机校准规范 JJG(纺织) 026 | (0～60)r/min | U=0.5r/min |  | 扩项 |
|  |  | 时间 |  |  | (0～30)min | U=0.20s |  |  |
|  |  | 温度 |  |  | (0～100)℃ | U=0.4℃ |  |  |
| 103 | *织物缩水率试验机 | 转速 | 700137 | 织物缩水率试验机校准规范 JJF(纺织) 052 | (0～60)r/min | U=0.5r/min |  | 扩项 |
|  |  | 时间 |  |  | (0～30)min | U=0.20s |  |  |
|  |  | 温度 |  |  | (0～100)℃ | U=0.4℃ |  |  |
| 104 | *汗渍色牢度仪 | 尺寸 | 700127 | 汗渍色牢度仪校准规范 JJF(纺织) 028 | (0～300)mm | U=0.03mm |  | 扩项 |
|  |  | 力值 |  |  | (0～100)N | Urel=0.4% |  |  |
| 105 | 数字多用表 | 直流电压 | 640504 | 直流数字电压表检定规程 JJG315,交流数字电压表检定规程 JJG(航天)34,(试行)直流数字电流检定规程 JJG598,交流数字电流表检定规程 JJG(航天)35,直流数字式欧姆表检定规程 JJG724 | 0.1mV～100mV | U=0.002%Vx+2μV |  |  |
|  |  |  |  |  | 100mV～1V | U=0.0011%Vx+3μV |  |  |
|  |  |  |  |  | 1V～10V | U=0.0012%Vx+15μV |  |  |
|  |  |  |  |  | 10V～100V | U=0.0018%VX+0.20mV |  |  |
|  |  |  |  |  | 100V～1000V | U=0.0018%Vx+1.8mV |  |  |
|  |  | 交流电压 |  |  | 1mV～100mV (45Hz～10kHz) | U=0.02%Vx+5μV |  |  |
|  |  |  |  |  | 100mV～1V (45Hz～10kHz) | U=0.018%Vx+5μV |  |  |
|  |  |  |  |  | 1V～10V (45Hz～10kHz) | U=0.012%Vx+15μV |  |  |
|  |  |  |  |  | 10V～100V (45Hz～1kHz) | U=0.018%Vx+0.5mV |  |  |
|  |  |  |  |  | 100V～1000V (45Hz～1kHz) | U=0.036%Vx+11mV |  |  |
|  |  | 直流电流 |  |  | (0.1～100)μA | U=0.003%Ix+0.002μA |  |  |
|  |  |  |  |  | 100μA～1mA | U=0.003%Ix+0.02μA |  |  |
|  |  |  |  |  | 1mA～10mA | U=0.003%Ix+0.2μA |  |  |
|  |  |  |  |  | 10mA～100mA | U=0.004%Ix+2μA |  |  |
|  |  |  |  |  | 100mA～1A | U=0.022%Ix+0.02mA |  |  |
|  |  |  |  |  | 1A～2.2A | U=0.035%IX+51uA |  |  |
|  |  |  |  |  | 2.2A～11A | U=0.07%IX+0.39mA |  |  |
|  |  | 交流电流 |  |  | 19μA～100μA (45Hz～1kHz) | U=0.08%Ix+0.04μA |  |  |
|  |  |  |  |  | 100μA～100mA (45Hz～1kHz) | U=0.08%Ix+(0.3μA～0.03mA) |  |  |
|  |  |  |  |  | 100mA～1A (45Hz～1kHz) | U=0.09%Ix+0.3mA |  |  |
|  |  |  |  |  | 1A～2.2A (45Hz～1kHz) | U=0.12%Ix+0.35mA |  |  |
|  |  |  |  |  | 2.2A～11A (45Hz～1kHz) | U=0.12%Ix+2.4mA |  |  |
|  |  | 电阻 |  |  | 0.1Ω～10Ω | U=0.003%Rx+0.0001Ω |  |  |
|  |  |  |  |  | 10Ω～100Ω | U=0.0026%Rx+0.0006Ω |  |  |
|  |  |  |  |  | 100Ω～1kΩ | U=0.002%Rx+0.01Ω |  |  |
|  |  |  |  |  | 1kΩ～10kΩ | U=0.002%Rx+0.06Ω |  |  |
|  |  |  |  |  | 10kΩ～100kΩ | U=0.002%x+0.6Ω |  |  |
|  |  |  |  |  | 100kΩ～1MΩ | U=0.003%Rx+3Ω |  |  |
|  |  |  |  |  | 1MΩ～10MΩ | U=0.006%Rx+12Ω |  |  |
|  |  |  |  |  | 10MΩ～100MΩ | U=0.06%x+0.5kΩ |  |  |
|  |  |  |  |  | 100MΩ～300MΩ | U=0.6%x+5kΩ |  |  |
| 106 | 多功能校准仪 | 直流电压 | 640503 | 直流标准电压源检定规程 JJG445,精密交流电压校准源检定规程 JJG410,直流标准电流源检定规程 JJG（航天）38,交流标准电流源检定规程 JJG(航天)51,直流电阻器检定规程 JJG166 | 10mV～100mV | U=1.1×10-5Vx+0.0004mV |  |  |
|  |  |  |  |  | 100mV～1V | U=1.0×10-5Vx+0.0004mV |  |  |
|  |  |  |  |  | 1V～10V | U=1.0×10-5Vx+0.0006mV |  |  |
|  |  |  |  |  | 10V～100V | U=1.2×10-5Vx+0.04mV |  |  |
|  |  |  |  |  | 100V～1000V | U=1.2×10-5Vx+0.2mV |  |  |
|  |  | 交流电压 |  |  | 10mV～10V (40Hz～1kHz) | U=8.1×10-5Vx+0.3mV |  |  |
|  |  |  |  |  | 10V～100V (40Hz～1kHz) | U=2.4×10-4Vx+0.003V |  |  |
|  |  |  |  |  | 100V～1000V (40Hz～1kHz) | U=4.7×10-4Vx+0.03V |  |  |
|  |  | 直流电流 |  |  | 100μA～1mA | U=2.3×10-5Ix+0.006μA |  |  |
|  |  |  |  |  | 1mA～10mA | U=2.3×10-5Ix+0.06μA |  |  |
|  |  |  |  |  | 10mA～100mA | U=4.1×10-5Ix+0.6μA |  |  |
|  |  |  |  |  | 100mA～1A | U=1.3×10-4Ix+0.02μA |  |  |
|  |  | 交流电流 |  |  | 1mA～100mA (45Hz～100Hz) | U=7.0×10-4Ix+0.03mA |  |  |
|  |  |  |  |  | 100mA～1A (45Hz～100Hz) | U=7.0×10-4Ix+0.3mA |  |  |
|  |  | 电阻 |  |  | 100Ω～1kΩ | U=1.2×10-5Rx+0.6mΩ |  |  |
|  |  |  |  |  | 1kΩ～10kΩ | U=1.2×10-5Rx+6mΩ |  |  |
|  |  |  |  |  | 10kΩ～100kΩ | U=1.2×10-5Rx+0.06Ω |  |  |
|  |  |  |  |  | 100kΩ～1MΩ | U=1.8×10-5Rx+3Ω |  |  |
|  |  |  |  |  | 1MΩ～10MΩ | U=5.8×10-5Rx+0.2kΩ |  |  |
| 107 | 交、直流电压电流表 | 直流电压 | 640599 | 电流表、电压表、功率表及电阻表检定规程 JJG124 | 0.3mV～330mV | Urel=0.13％ |  |  |
|  |  |  |  |  | 0.33V～1000V | Urel=0.12％ |  |  |
|  |  | 交流电压 |  |  | 0.3mV～330mV (45Hz～1kHz) | Urel=0.16％ |  |  |
|  |  |  |  |  | 0.33V～33V (45Hz～1kHz) | Urel=0.13％ |  |  |
|  |  |  |  |  | 33V～1000V (45Hz～1kHz) | Urel=0.14％ |  |  |
|  |  | 直流电流 |  |  | 0.01mA～3.3mA | Urel=0.15％ |  |  |
|  |  |  |  |  | 3.3mA～330mA | Urel=0.14％ |  |  |
|  |  |  |  |  | 0.33A～11A | Urel=0.13％ |  |  |
|  |  | 交流电流 |  |  | 0.03mA～3.3mA (45Hz～1kHz) | Urel=0.23％ |  |  |
|  |  |  |  |  | 3.3mA～33mA (45Hz～1kHz) | Urel=0.20％ |  |  |
|  |  |  |  |  | 0.033A～2.2A (45Hz～1kHz) | Urel=0.19％ |  |  |
|  |  |  |  |  | 2.2A～11A (45Hz～1kHz) | Urel=0.18％ |  |  |
| 108 | 直流稳压电源 | 直流电压 | 640103 | 直流稳压电源检定规程 JJG(航天)6 | 0.2V～1V | Urel=0.019% |  |  |
|  |  |  |  |  | 1V～10V | Urel=0.015% |  |  |
|  |  |  |  |  | 10V～100V | Urel=0.018% |  |  |
|  |  |  |  |  | 100V～400V | Urel=0.021% |  |  |
|  |  | 直流电流 |  |  | 0.1A～0.2A | Urel=0.052% |  |  |
|  |  |  |  |  | 0.2A～2A | Urel=0.038% |  |  |
|  |  |  |  |  | 2A～30A | Urel=0.037% |  |  |
| 109 | 交流稳压电源(变频电源) | 交流电压 | 640203 | 精密交流电压校准源检定规程 JJG410 | 1V～10V (50Hz～1kHz) | Urel=0.054% |  |  |
|  |  |  |  |  | 10V～100V (50Hz～1kHz) | Urel=0.059% |  |  |
|  |  |  |  |  | 100V～750V (50Hz～1kHz) | Urel=0.063% |  |  |
|  |  | 交流电流 |  |  | 0.1A～0.2A (50Hz～400Hz) | Urel=0.13% |  |  |
|  |  |  |  |  | 0.2A～2A (50Hz～400Hz) | Urel=0.097% |  |  |
|  |  |  |  |  | 2A～20A (50Hz～400Hz) | Urel=0.11% |  |  |
|  |  | 频率 |  |  | (50Hz～1kHz) | Urel=0.002% |  |  |
| 110 | 直流数字欧姆表 | 电阻 | 640305 | 直流数字式欧姆表检定规程 JJG724 | 0.01Ω | Urel=0.023% |  |  |
|  |  |  |  |  | 0.1Ω | Urel=0.024% |  |  |
|  |  |  |  |  | 0.1Ω～1Ω | Urel=0.042% |  |  |
|  |  |  |  |  | 1Ω～10Ω | Urel=0.035% |  |  |
|  |  |  |  |  | 10Ω～100Ω | Urel=0.034% |  |  |
|  |  |  |  |  | 100Ω～1000Ω | Urel=0.032% |  |  |
|  |  |  |  |  | 1kΩ～10kΩ | Urel=0.035% |  |  |
|  |  |  |  |  | 10k～100kΩ | Urel=0.076% |  |  |
|  |  |  |  |  | 0.1MΩ～1MΩ | Urel=0.083% |  |  |
| 111 | 直流低电阻表 | 电阻 | 640306 | 直流低电阻表检定规程 JJG837 | 0.001Ω | Urel=0.023% |  |  |
|  |  |  |  |  | 0.01Ω | Urel=0.024% |  |  |
|  |  |  |  |  | 0.1Ω～1Ω | Urel=0.042% |  |  |
|  |  |  |  |  | 1Ω～10Ω | Urel=0.035% |  |  |
|  |  |  |  |  | 10Ω～100Ω | Urel=0.034% |  |  |
|  |  |  |  |  | 100Ω～1000Ω | Urel=0.032% |  |  |
|  |  |  |  |  | 1kΩ～10kΩ | Urel=0.035% |  |  |
|  |  |  |  |  | 10k～100kΩ | Urel=0.076% |  |  |
|  |  |  |  |  | 0.1MΩ～1MΩ | Urel=0.083% |  |  |
| 112 | 直流电阻箱 | 电阻 | 640304 | 直流电阻箱检定规程 JJG982 | 0.1Ω～1Ω | Urel=0.056% |  |  |
|  |  |  |  |  | 1Ω～10Ω | Urel=0.019% |  |  |
|  |  |  |  |  | 10Ω～100Ω | Urel=0.017% |  |  |
|  |  |  |  |  | 100Ω～1000Ω | Urel=0.014% |  |  |
|  |  |  |  |  | 1kΩ～10MΩ | Urel=0.013% |  |  |
| 113 | 数字功率计 | 电压 | 640401 | 交流数字功率表检定规程 JJG780 | 0.1V～3.3V (45Hz～1kHz) | Urel=0.12% |  |  |
|  |  |  |  |  | 3.3V～1000V (45Hz～1kHz) | Urel=0.11% |  |  |
|  |  | 电流 |  |  | 0.1A～2.2A (45Hz～1kHz) | Urel=0.13% |  |  |
|  |  |  |  |  | 2.2A～11A (45Hz～1kHz) | Urel=0.11% |  |  |
|  |  | 功率 |  |  | 10mW～330mW (45Hz～65Hz) | Urel=0.14% |  |  |
|  |  |  |  |  | 330mW～3.3W (45Hz～65Hz) | Urel=0.13% |  |  |
|  |  |  |  |  | 3.3W～3.3kW (45Hz～65Hz) | Urel=0.12% |  |  |
|  |  |  |  |  | 3.3kW～11kW (45Hz～65Hz) | Urel=0.13% |  |  |
| 114 | 直流电桥 | 电阻 | 640307 | 直流电桥检定规程 JJG125 | 1Ω～10Ω | Urel=0.25% |  |  |
|  |  |  |  |  | 10Ω～100Ω | Urel=0.06% |  |  |
|  |  |  |  |  | 100Ω～10kΩ | Urel=0.02% |  |  |
| 115 | 高阻表 | 电阻 | 640309 | 高绝缘电阻测量仪检定规程 JJG690 | 100Ω～10MΩ | Urel=0.7% |  |  |
|  |  |  |  |  | 10MΩ～100MΩ | Urel=0.8% |  |  |
|  |  |  |  |  | 100MΩ～1GΩ | Urel=1.2% |  |  |
|  |  |  |  |  | 1GΩ～10GΩ | Urel=2.6% |  |  |
|  |  |  |  |  | 10GΩ～100GΩ | Urel=6.0% |  |  |
| 116 | 接地导通电阻测试仪 | 电阻 | 640310 | 接地导通电阻测试仪检定规程 JJG984 | 20mΩ～120mΩ | Urel=0.3% |  |  |
|  |  | 电流 |  |  | 1A～30A | Urel=0.6% |  |  |
| 117 | 绝缘电阻表(兆欧表) | 电阻 | 640308 | 绝缘电阻表（兆欧表）检定规程 JJG622 | 100Ω～10MΩ | Urel=1.0% |  |  |
|  |  |  |  |  | 10MΩ～100MΩ | Urel=1.2% |  |  |
|  |  |  |  |  | 100MΩ～1GΩ | Urel=1.6% |  |  |
|  |  |  |  |  | 1GΩ～10GΩ | Urel=2.6% |  |  |
|  |  | 电压 |  |  | 0.1kV～10kV | Urel=0.6% |  |  |
| 118 | 漏电流测试仪 | 电压 | 640206 | 泄漏电流测试仪检定规程 JJG843 | 50V～650V | Urel=0.11% |  |  |
|  |  | 电流 |  |  | 10μA～100mA | Urel=0.2% |  |  |
| 119 | 电池内阻测试仪 | 电阻 | 711015 | 直流数字电压表检定规程 JJG315,直流低电阻表检定规程 JJG837 | 10mΩ～10Ω | Urel=0.25% |  |  |
|  |  |  |  |  | 10Ω～100Ω | Urel=0.06% |  |  |
|  |  | 直流电压 |  |  | 0.1V～100 V | Urel=0.15% |  |  |
| 120 | 静电手环(带)/脚环(带)测试仪(静电鞋测试仪) | 电阻 | 640312 | 电子式绝缘电阻表检定规程 JJG1005,静电腕带/脚盘测试仪校准规范 JJF（电子）31502 | 100Ω～10MΩ | Urel=0.8% |  |  |
|  |  |  |  |  | 10MΩ～100MΩ | Urel=1.2% |  |  |
| 121 | 电子式绝缘电阻表 | 电阻 | 640308 | 电子式绝缘电阻表检定规程 JJG1005 | 100Ω～10MΩ | Urel=0.7% |  |  |
|  |  |  |  |  | 10MΩ～100MΩ | Urel=0.9% |  |  |
|  |  |  |  |  | 100MΩ～1GΩ | Urel=1.2% |  |  |
|  |  |  |  |  | 1GΩ～10GΩ | Urel=2.6% |  |  |
|  |  |  |  |  | 10GΩ～100GΩ | Urel=6% |  |  |
|  |  | 电压 |  |  | 0.1kV～10kV | Urel=0.6% |  |  |
| 122 | 安规测试仪 | 直流电压 | 640599 | 耐电压测试仪检定规程 JJG795,电子式绝缘电阻表检定规程 JJG1005,泄漏电流测量仪检定规程 JJG843,接地导通电阻测试仪检定规程 JJG984 | 0.5kV～10kV | Urel=1.0% |  |  |
|  |  | 交流电压 |  |  | 0.5kV～10kV (45Hz～65Hz) | Urel=1.8% |  |  |
|  |  | 直流电流 |  |  | 0.5mA～100mA | Urel=0.7% |  |  |
|  |  | 交流电流 |  |  | 0.5mA～100mA(45Hz～65Hz) | Urel=1.2% |  |  |
|  |  | 时间 |  |  | 10s～999.9s | Urel=0.7% |  |  |
|  |  | 电阻 |  |  | 100Ω～10MΩ | Urel=0.7% |  |  |
|  |  |  |  |  | 10MΩ～100MΩ | Urel=0.9% |  |  |
|  |  |  |  |  | 100MΩ～1GΩ | Urel=1.2% |  |  |
|  |  |  |  |  | 1GΩ～10GΩ | Urel=2.6% |  |  |
|  |  |  |  |  | 10GΩ～100GΩ | Urel=6% |  |  |
|  |  | 电压 |  |  | 0.1kV～1kV | Urel=0.6% |  |  |
|  |  | 交流电压 |  |  | 50V～650V (45Hz～65Hz) | Urel=0.11% |  |  |
|  |  | 交流电流 |  |  | 10μA～100mA (45Hz～65Hz) | Urel=0.2% |  |  |
|  |  | 电阻 |  |  | 20mΩ～120mΩ | Urel=0.3% |  |  |
|  |  | 电流 |  |  | 1A～30A | Urel=0.6% |  |  |
| 123 | 耐电压测试仪 | 直流电压 | 640113 | 耐电压测试仪检定规程 JJG795 | 0.5kV～10kV | Urel=1.0% |  |  |
|  |  | 交流电压 |  |  | 0.5kV～10kV (45Hz～65Hz) | Urel=1.8% |  |  |
|  |  | 直流电流 |  |  | 0.1mA～100mA | Urel=0.7% |  |  |
|  |  | 交流电流 |  |  | 0.1mA～100mA(45Hz～65Hz) | Urel=1.2% |  |  |
|  |  | 时间 |  |  | 10s～999.9s | Urel=0.7% |  |  |
| 124 | 线材测试仪 | 耐压(交流） | 701018 | 线缆测试仪校准规范 JJF1457 | (10～1500)V | Urel=1.8% |  |  |
|  |  | 耐压(直流） |  |  | (10～1500)V | Urel=1.0% |  |  |
|  |  | 泄漏电流(交流) |  |  | (0～10)mA | Urel=1.2% |  |  |
|  |  | 泄漏电流(直流) |  |  | (0～10)mA | Urel=0.7% |  |  |
|  |  | 导通电阻 |  |  | 0Ω～100kΩ | Urel=1.2% |  |  |
|  |  | 绝缘电阻 |  |  | 200kΩ～1000MΩ | Urel=1.8% |  |  |
|  |  | 直流电压 |  |  | 10V～1000V | Urel=0.3% |  |  |
| 125 | *电线电缆火花试验机 | 直流电压 | 711014 | 火花试验机校准规范 JJG（浙）84 | 0.5kV～10kV | Urel=0.7% |  |  |
|  |  | 交流电压 |  |  | 0.5kV～10kV (45～65)Hz | Urel=1.3% |  |  |
| 126 | 电子负载 | 直流电压 | 701018 | 直流电子负载校准规范 JJF1462 | 0.33V～33V | U=5.8×10-5Vx+58μV |  |  |
|  |  |  |  |  | 33V～250V | U=6.4×10-5Vx+0.58mV |  |  |
|  |  | 直流电流 |  |  | 0.1A～10A | Urel=0.17% |  |  |
|  |  |  |  |  | (10～200)A | Urel=0.26% |  |  |
| 127 | 钳形电流表 | 交流电流 | 640204 | 钳形电流表校准规范 JJF1075 | 0.33A～2.2A (45Hz～1kHz) | U=1.2×10-3Ix+2.3mA |  | 变更 |
|  |  |  |  |  | 2.2A～11A (45Hz～1kHz) | U=1.2×10-3Ix+0.35mA |  |  |
|  |  |  |  |  | 11A～1000A (45Hz～1kHz) | Urel=0.35% |  |  |
|  |  | 直流电流 |  |  | 0.33A～2.2A | U=3.5×10-4Ix+51μA |  |  |
|  |  |  |  |  | 2.2A～11A | U=7.0×10-4Ix+0.39mA |  |  |
|  |  |  |  |  | 11A～1000A | Urel=0.35% |  |  |
| 128 | 接地电阻表 | 电阻 | 640310 | 接地电阻表检定规程 JJG366 | 0.01Ω～0.1Ω | Urel=0.43% |  |  |
|  |  |  |  |  | 0.1Ω～1Ω | Urel=0.35% |  |  |
|  |  |  |  |  | 1Ω～10Ω | Urel=0.27% |  |  |
|  |  |  |  |  | 10Ω～100Ω | Urel=0.25% |  |  |
|  |  |  |  |  | 100Ω～1000Ω | Urel=0.23% |  |  |
| 129 | 表面电阻测试仪 | 电阻 | 640312 | 表面电阻测试仪校准规范 JJF1285 | 100Ω～10MΩ | Urel=0.7% |  |  |
|  |  |  |  |  | 10MΩ～100MΩ | Urel=0.8% |  |  |
|  |  |  |  |  | 100MΩ～1GΩ | Urel=1.2% |  |  |
|  |  |  |  |  | 1GΩ～10GΩ | Urel=2.6% |  |  |
|  |  |  |  |  | 10GΩ～100GΩ | Urel=6.0% |  |  |
|  |  | 电压 |  |  | 1V～250V | Urel=0.3% |  |  |
| 130 | 数字高压表 | 交流电压 | 640199 | 数字高压表校准规范 JJF(鲁)60 | (0.1～10.0)kV | Urel=0.8% |  | 扩项 |
|  |  | 直流电压 |  |  | (0.1～10.0)kV | Urel=1.5% |  |  |
| 131 | 高压静电电压表 | 交流电压 | 640199 | 高压静电电压表检定规程 JJG 494 | (0.6～10.0)kV | Urel=1.2% |  | 扩项 |
|  |  | 直流电压 |  |  | (0.6～10.0)kV | Urel=1.2% |  |  |
| 132 | 工频高压分压器 | 分压比 | 640107 | 工频高压分压器检定规程 JJG 496 | (1～10.0)kV | Urel=0.5% |  | 扩项 |
| 133 | 直流电阻器 | 电阻 | 640303 | 直流电阻器检定规程 JJG166 | 1mΩ～100kΩ | Urel=0.1%~0.012% |  | 扩项 |
| 134 | 直流标准电压源 | 直流电压 | 640101 | 直流标准电压源检定规程 JJG445 | 10mV～1000V | Urel=0.02% |  | 扩项 |
| 135 | 直流标准电流源 | 直流电流 | 640201 | 直流标准电流源检定规程 JJG（航天）38 | (0～200)A | Urel=0.5% |  | 扩项 |
| 136 | 交流标准电压源 | 交流电压 | 640102 | 交流标准电压源检定规程 JJG(航天)2 | 10mV～1000V | 45Hz～10kHz:Urel=0.05% |  | 扩项 |
| 137 | 交流标准电流源 | 交流电流 | 640202 | 交流标准电流源检定规程 JJG(航天)51 | 0.2mA～100A | Urel=0.8% |  | 扩项 |
| 138 | 带电绕组温升测量仪 | 电阻 | 640399 | 直流低电阻表检定规程 JJG837 | 0.01Ω～1kΩ | Urel=0.2％ |  | 扩项 |
| 139 | 相位表(功率因数表) | 功率因数 | 640406 | 工频单相相位表检定规程 JJG 440 | -1～+1 | U=0.002 |  | 扩项 |
| 140 | 直流电阻分压箱 | 分压比 | 640107 | 直流电阻分压箱检定规程 JJG531 | (10～500):1 | Urel=0.04% |  | 扩项 |
| 141 | *插头综合测试仪 | 绝缘电阻 | 640599 | 耐电压测试仪检定规程 JJG795,电子式绝缘电阻表检定规程 JJG1005 | (1～1000)MΩ | Urel=1.2% |  | 扩项 |
|  |  | 交流电压 |  |  | (0.5～10.0)kV | Urel=1.2% |  |  |
|  |  | 交流电流 |  |  | (0.1～50)mA | Urel=1.2% |  |  |
|  |  | 时间 |  |  | (1～1000)s | Urel=0.9% |  |  |
| 142 | 电容漏电流测量仪 | 直流电压 | 640206 | 电容器漏电流测试仪 JJ（电子）306003 | (1～1000)V | Urel=0.16% |  | 扩项 |
|  |  | 直流电流 |  |  | (0.001～30)mA | Urel=0.4% |  |  |
| 143 | *针焰试验仪 | 长度 | 701050 | 电工电子产品着火危险试验第5部分：试验火焰针焰试验 GB/T5169.5 | (0.1～100)mm | U=0.03mm |  | 扩项 |
|  |  | 角度 |  |  | (0～90)° | U=0.3° |  |  |
|  |  | 温度 |  |  | (100～1000)℃ | U=1.5℃ |  |  |
|  |  | 时间 |  |  | (1～3600)s | Urel=1.2% |  |  |
| 144 | 电子式时间继电器 | 时间 | 660213 | 电子式时间继电器校准规范 JJF1282 | (1～3600)s | Urel=5x10-5s |  | 扩项 |
| 145 | 直流分流器 | 电阻 | 640314 | 直流分流器检定规程 JJG 1069 | (0.1～1000)mΩ | Urel=0.25% |  | 扩项 |
| 146 | 高频火花机 | 交流电压 | 701014 | 火花试验机检定规程 JJG（苏）74 | (0.5～15)kV | Urel=1.8% |  | 扩项 |
|  |  | 频率 |  |  | 50Hz～10kHz | Urel=2% |  |  |
|  |  | 击穿电流 |  |  | (0.1～0.6)mA | Urel=1.2% |  |  |
| 147 | 电容表 | 电容 | 650311 | 交流电桥检定规程 JJG 441 | 100pF～1µF | Urel=0.3％ |  | 扩项 |
|  |  |  |  |  | 10nF～1000µF | Urel=0.4%~0.6％ |  |  |
| 148 | 线圈圈数测量仪 | 圈数 | 701051 | 线圈圈数测量仪校准规范 JJF（浙）1065 | (1～10000)T | Urel=0.2% |  | 扩项 |
|  |  |  |  |  | (10000～20000)T | Urel=0.25% |  |  |
| 149 | *冲击全波电压试验装置 | 电压 | 640199 | 20kV及以下冲击全波电压试验装置校准规范 JJF（机械）020 | (0.5～20)kV | Urel=3% |  | 扩项 |
|  |  | 上升时间 |  |  | (0～5)ms | Urel=5% |  |  |
|  |  | 半峰时间 |  |  | (0～5)ms | Urel=5% |  |  |
| 150 | *冲击振荡电流干扰源 | 交流电流 | 640299 | 冲击振荡电流干扰源校准规范 JJF（机械）023 | (0～500)A | Urel=3% |  | 扩项 |
|  |  | 视在波前时间 |  |  | (0～5)ms | Urel=5% |  |  |
|  |  | 衰减率 |  |  | (0～100)% | Urel=6% |  |  |
| 151 | 匝间绝缘耐压试验机 | 峰值电压 | 701052 | 耐电压测试仪检定规程 JJG795 | (0.5～5)kV | Urel=2.5% |  | 扩项 |
| 152 | *漏电起痕试验仪 | 电压 | 701016 | 漏电起痕试验仪校准规范 JJF(浙) 1087 | (10～1000)V | Urel=0.5% |  | 扩项 |
|  |  | 电流 |  |  | (0.1～10)A | Urel=0.6% |  |  |
|  |  | 时间 |  |  | (1～3600)s | U=0.26s |  |  |
|  |  | 长度 |  |  | (0.5～150)mm | U=0.03mm |  |  |
|  |  | 角度 |  |  | (15.1～90)° | U=0.5° |  |  |
|  |  | 力值 |  |  | (0.5～5)N | Urel=1.2% |  |  |
| 153 | 接地电阻测量仪检定装置 | 电阻 | 640599 | 直流电阻箱检定规程 JJG982 | (10～500)mΩ | Urel=0.17% |  | 扩项 |
| 154 | 钳形接地电阻仪 | 电阻误差 | 640310 | 钳形接地电阻仪 JJG 1054 | (0.01～0.1)Ω | Urel=8% |  | 扩项 |
|  |  |  |  |  | (0.1～1) Ω | Urel=2.4% |  |  |
|  |  |  |  |  | (1～10) Ω | Urel=0.4% |  |  |
|  |  |  |  |  | (10～10000) Ω | Urel=0.3% |  |  |
| 155 | *静电放电发生器 | 电压 | 701028 | 静电放电发生器校准规范 JJF(电子) 30801 | (0.5～20)kV | Urel=3% |  | 扩项 |
|  |  | 上升时间 |  |  | 1ns～10ms | Urel=5% |  |  |
| 156 | *电浪涌发生器 | 电压 | 701030 | 电浪涌发生器校准规范 JJF(电子)30803 | (0.5～10)kV | Urel=3% |  | 扩项 |
|  |  | 时间 |  |  | 1ns～10ms | Urel=5% |  |  |
| 157 | *电快速瞬变脉冲群发生器 | 电压 | 701029 | 电快速瞬变脉冲群发生器校准规范 JJF（电子）30804 | 100V～8kV | Urel=3% |  | 扩项 |
|  |  | 时间 |  |  | 1ns～100ms | Urel=3% |  |  |
| 158 | *钢筋锈蚀测量仪 | 输出电压 | 700840 | 钢筋锈蝕测量仪校准规范 JJF1341 | 20mV～20V | Urel=0.26% |  | 扩项 |
|  |  | 输出电流 |  |  | 0.01mA～2A | Urel=0.29% |  |  |
|  |  | 电位测量 |  |  | 0～10V | Urel=0.32% |  |  |
| 159 | 晶体管特性图示仪 | 电压 | 650501 | 半导体管特性图示仪校准规范 JJF1236 | 40mV～330mV | U=7.0×10-5Vx+4μV |  |  |
|  |  |  |  |  | 0.33V～3.3V | U=5.8×10-5Vx+6μV |  |  |
|  |  |  |  |  | 3.3V～33V | U=5.8×10-5Vx+58μV |  |  |
|  |  |  |  |  | 33V～330V | U=6.4×10-5Vx+578μV |  |  |
|  |  |  |  |  | 330V～1000V | U=7.0×10-5Vx+1732μV |  |  |
|  |  | 电流 |  |  | 0.4mA～3.3mA | U=1.5×10-4Ix+0.06μA |  |  |
|  |  |  |  |  | 3.3mA～33mA | U=1.2×10-4Ix+0.29μA |  |  |
|  |  |  |  |  | 33mA～330mA | U=1.2×10-4Ix+3.9μA |  |  |
|  |  |  |  |  | 0.33A～2.2A | U=3.5×10-4Ix+51μA |  |  |
|  |  |  |  |  | 2.2A～10A | U=7.0×10-4Ix+382μA |  |  |
|  |  | 阶梯电压 |  |  | 1mV～400V | Urel=4% |  |  |
|  |  | 阶梯电流 |  |  | 1μA～10A | Urel=4% |  |  |
| 160 | 低频电子电压表 | 交流电压 | 650105 | 低频电子电压表检定规程 JJG782 | 1mV～33mV (10Hz～100kHz) | Urel=1.2% |  |  |
|  |  |  |  |  | 33mV～330mV (10Hz～100kHz) | Urel=0.96% |  |  |
|  |  |  |  |  | 330mV～3.3V (10Hz～100kHz) | Urel=0.65% |  |  |
|  |  |  |  |  | 3.3V～33V (10Hz～100kHz) | Urel=0.63% |  |  |
|  |  |  |  |  | 33V～300V (10Hz～100kHz) | Urel=0.60% |  |  |
| 161 | 低频信号发生器 | 频率 | 650205 | 低频信号发生器检定规程 JJG602 | 10Hz～100kHz | Urel=0.01% |  |  |
|  |  | 电压 |  |  | 1mV～20V | Urel=1.5% |  |  |
|  |  | 失真度 |  |  | (0.03～30)% | Urel=2% |  |  |
| 162 | 音频(扫频)信号发生器 | 频率 | 650206 | 声频信号发生器检定规程 JJG607 | 10Hz～200kHz | Urel=1×10-7 |  |  |
|  |  | 电压 |  |  | 1mV～20V | Urel=1.5% |  |  |
|  |  | 失真度 |  |  | (0.03～30)% | Urel=2% |  |  |
| 163 | 函数信号发生器 | 频率 | 650201 | 函数信号发生器检定规程 JJG840 | 1Hz～30MHz | Urel=4×10-7 |  |  |
|  |  | 电压 |  |  | 1mV～10V (1Hz～30MHz) | Urel=3.5% |  |  |
|  |  | 调幅 |  |  | (0～99)% | Urel=1.6% |  |  |
|  |  | 调频 |  |  | (0～100)kHz | Urel=1.4% |  |  |
|  |  | 失真度 |  |  | (0.03～30)% | Urel=2% |  |  |
| 164 | 高频信号 发生器 | 频率 | 650207 | 信号发生器检定规程 JJG173 | 0.1MHz～18GHz | Urel=2×10-7 |  |  |
|  |  | 电平 |  |  | -140dBm～0dBm(0.1MHz～1040MHz) | U=0.2dB |  |  |
|  |  |  |  |  | 0dBm～17dBm(0.1MHz～1040MHz) | U=0.4dB |  |  |
|  |  | 调幅 |  |  | (0～99)% | Urel=1.6% |  |  |
|  |  | 调频 |  |  | (0～400)kHz | Urel=1.4% |  |  |
| 165 | FO测试仪 | 频率 | 650299 | 声频信号发生器检定规程 JJG607 | 16Hz～3kHz | Urel=0.2% |  |  |
| 166 | 示波器 | 垂直偏转系数 | 650212 | 模拟示波器检定规程 JJG262 | (10μV～20V)/div | Urel=1.1％ |  |  |
|  |  | 扫描时间系数 |  |  | (2ns～5s)/div | Urel=0.7％ |  |  |
|  |  | 带宽 |  |  | ≤200MHz | Urel=3.2％ |  |  |
| 167 | 示波器电压探头 | 直流电压衰减比 | 701044 | 示波器电压探头校准规范 JJF 1437 | 1:1～1:1000 | Urel=0.70% |  | 扩项 |
|  |  | 频率响应 |  |  | 50kHz～200MHz | U=0.72dB |  |  |
|  |  | 上升时间 |  |  | 700ps～10ms | Urel=6% |  |  |
| 168 | 示波器电流探头 | 直流衰减系数 | 701044 | 示波器电流探头校准规范 JJF（电子）30305 | 1～100 | Urel=1.2% |  | 扩项 |
|  |  | 交/直流电流 |  |  | ≤1000A | Urel=1.0% |  |  |
|  |  | 频带宽度 |  |  | ≤200MHz | Urel=4.7% |  |  |
|  |  | 上升时间 |  |  | ≥0.7ns | U=0.05ns |  |  |
| 169 | 示波器差分探头 | 共模抑制比 | 701044 | 示波器差分探头校准规范 JJF（电子）30306 | ≥18dB | U=1.6dB |  | 扩项 |
|  |  | 直流衰减系数 |  |  | 1～1000 | Urel=1.2% |  |  |
|  |  | 频带宽度 |  |  | 50kHz～200MHz | Urel=4.7% |  |  |
|  |  | 上升时间 |  |  | ≥0.7ns | U=0.05ns |  |  |
|  |  | 输入电阻 |  |  | (0～100)MΩ | Urel=0.1% |  |  |
| 170 | 驻极体传声器测试仪 | 频率 | 630211 | 驻极体传声器测试仪校准规范 JJF1145 | 20Hz～100kHz | Urel=0.003% |  | 扩项 |
|  |  | 电压 |  |  | (0.1～10)V | Urel=0.1% |  |  |
|  |  | 电流 |  |  | (0.1～10)mA | Urel=0.4% |  |  |
| 171 | 电话分析仪 | 电平 | 650610 | 双音多频电话机测试器检定规程 JJG(YD)032 | (-30～0)dB | U=0.18dB |  | 扩项 |
|  |  | 频率 |  |  | 10Hz～20kHz | Urel=0.02% |  |  |
| 172 | 频谱分析仪 | 频率 | 650124 | 频谱分析仪校准规范 JJF 1396 | 10Hz～4GHz | Urel=1×10-7 |  | 扩项 |
|  |  | 电平 |  |  | (-130～20)dBm | U=0.22dB |  |  |
| 173 | 合成信号发生器 | 频率 | 650299 | 合成信号发生器检定规程 JJG 502 | 5kHz～20 GHz | Urel=1×10-8 |  | 扩项 |
|  |  | 电平 |  |  | (-130～20)dBm | Urel=0.26dB |  |  |
|  |  | 调频 |  |  | 10Hz～400kHz | Urel=1.5% |  |  |
|  |  | 调幅 |  |  | 5%～99% | Urel=1.5% |  |  |
|  |  | 失真度 |  |  | 0.01%～100% | Urel=15% |  |  |
| 174 | *脉冲信号发生器 | 脉冲输出幅度 | 650204 | 脉冲信号发生器检定规程 JJG 490 | 10mV～200V | Urel=1.2% |  | 扩项 |
|  |  | 脉冲上升沿 |  |  | 1ns～1s | Urel=5% |  |  |
|  |  | 脉冲宽度 |  |  | 1ns～50ms | Urel=0.6% |  |  |
|  |  | 输出频率 |  |  | 1Hz～500MHz | Urel=1×10-8 |  |  |
| 175 | 电子电压表(高频毫伏表) | 交流电压 | 650105 | 电子电压表检定规程 JJG 250 | 1mV～330V(1kHz) | Urel=(0.8～0.3)% |  | 扩项 |
|  |  |  |  |  | 1mV～3.3V(1MHz) | Urel=1.3% |  |  |
|  |  | 频率响应 |  |  | 10Hz～125MHz | Urel=1.3% |  |  |
|  |  |  |  |  | 125Hz～500MHz | Urel=3.6% |  |  |
| 176 | 网络分析仪 | 频率 | 650125 | 网络分析仪 JJF（电子）30501 | 10kHz～20GHz | Urel=1×10-8 |  | 扩项 |
|  |  | 功率电平 |  |  | (-130～20)dBm | U=0.22dB |  |  |
|  |  | 传输幅度特性 |  |  | (0～-50)dB | U=0.20dB |  |  |
|  |  | 反射特性 |  |  | SWR:0～1.0（≤4GHz） | U=0.05 |  |  |
| 177 | 调制度测量仪 | 调幅 | 650208 | 调制度测量仪校准规范 JJF 1111 | 5%～99% | Urel=5% |  | 扩项 |
|  |  | 调频 |  |  | 10Hz～400kHz | Urel=5% |  |  |
| 178 | 电视视频信号发生器 | 电平 | 650601 | 电视视频信号发生器校准规范 JJF1235 | 1mV～1000mV | Urel=1.1% |  | 扩项 |
|  |  | 时间 |  |  | 1ns～1s | U=0.0025×R＋0.0114μs |  |  |
| 179 | *电视信号场强仪 | 频率 | 650602 | 电视信号场强仪检定规程 JJG1057 | (48～862)MHz | Urel=1×10-8 |  | 扩项 |
|  |  | 电平 |  |  | (30～120)dBμV | U=0.8dB |  |  |
| 180 | 示波器校准仪 | 直流电压 | 650213 | 示波器校准仪检定规程 JJG 278 | 1mV～1000V | Urel=0.04% |  | 扩项 |
|  |  | 方波电压 |  |  | (10Hz～10kHz)：1mV～100mV | U=0.09%Vx+0.06mV |  |  |
|  |  |  |  |  | (10Hz～10kHz)：100mV～1V | U=0.09%Vx+0.5mV |  |  |
|  |  |  |  |  | (10Hz～10kHz)：1V～10V | U=0.09%Vx+5mV |  |  |
|  |  |  |  |  | (10Hz～10kHz)：10V～100V | U=0.09%Vx+0.05V |  |  |
|  |  |  |  |  | (10Hz～10kHz)：100V～200V | U=0.09%Vx+0.5V |  |  |
|  |  | 时标 |  |  | 0.5ns～5s | Urel=1×10-7 |  |  |
|  |  | 上升时间 |  |  | 1ns～1s | Urel=5% |  |  |
| 181 | *蓝牙测试仪 | 频率 | 650611 | 蓝牙测试仪校准规范 JJF 1278 | (2.402～2.480)GHz | Urel=1×10-8 |  | 扩项 |
|  |  | 电平 |  |  | (-90～0)dBm | U=0.22dB |  |  |
|  |  | 谐波 |  |  | (0～-80)dBc | U=3.0dB |  |  |
|  |  | 非谐波 |  |  | (0～-80)dBc | U=3.0dB |  |  |
|  |  | GFSK调频指数 |  |  | 0.25～0.38 | Urel=0.7% |  |  |
| 182 | EMI测量接收机 | 频率 | 650122 | 电磁骚扰测量接收机校准规范 JJF1144 | 9 kHz～3 GHz | Urel=1×10-7 |  | 扩项 |
|  |  | 电平 |  |  | (-120～20)dBm | U=0.3dB |  |  |
| 183 | 矢量信号分析仪 | 频率 | 650216 | 矢量信号分析仪校准规范 JJF 1128 | 9kHz～4GHz | Urel=3×10-8 |  | 扩项 |
|  |  | 电平 |  |  | (-120～20)dBm | U = 0.4dB |  |  |
|  |  | EVM |  |  | (0.3～5)% | Urel=0.5% |  |  |
|  |  | 相位 |  |  | (0.17～5)° | U=0.24° |  |  |
|  |  | FSK |  |  | (0.5～5)% | Urel=1.7% |  |  |
| 184 | 功率传感器 | 校准因子 | 650110 | 脉冲功率计检定规程 JJG 1024 | 100kHz～ 4GHz | 10μW～1mW Urel=5.0% | 只用交替比较法 | 扩项 |
| 185 | 选频电平表 | 电平 | 650101 | 选频电平表检定规程 JJG777 | -100dB～20dB | U=0.2dB |  | 扩项 |
|  |  | 频率 |  |  | 20Hz～30MHz | Urel=1×10-6 |  |  |
|  |  | 频率响应 |  |  | -100dB～20dB | U =0.2dB |  |  |
| 186 | 射频通信测试仪 | 输出频率/电平 | 650607 | 射频通信测试仪校准规范 JJF1065 | 1Hz～4.2GHz | Urel=1×10-8 |  | 扩项 |
|  |  |  |  |  | (-130～20)dBm | U=0.22dB |  |  |
|  |  | 调幅 |  |  | 5%～99% | Urel=1.8% |  |  |
|  |  | 调相 |  |  | PM:≤400rad | Urel=1.2% |  |  |
|  |  | 频率测量 |  |  | 100kHz～4GHz | Urel=1×10-7 |  |  |
|  |  | 功率测量 |  |  | (-130～20)dBm | U=0.22dB |  |  |
|  |  | 调幅测量 |  |  | 5%～99% | Urel=1.8% |  |  |
|  |  | 调频测量 |  |  | 10Hz～400kHz | Urel=1.8% |  |  |
|  |  | 调相测量 |  |  | (0～100)rad | Urel=1.2% |  |  |
|  |  | 音频频率 |  |  | 20Hz～200kHz | Urel=0.01% |  |  |
|  |  | 电压 |  |  | 1mV～300V | Urel=0.5% |  |  |
|  |  | 失真度 |  |  | (0.05～30)% | Urel=3.6% |  |  |
| 187 | 信纳表 | 电平 | 630207 | 信纳表校准规范 JJF 1165 | (0～20)dB | U=0.2dB |  | 扩项 |
| 188 | 失真度测量仪 | 失真度 | 650209 | 失真度测量仪检定规程 JJG 251 | 0.01% ～30% | Urel=0.3% |  | 扩项 |
|  |  | 电压 |  |  | 1mV～300V | Urel=0.3% |  |  |
| 189 | 音频分析仪 | 失真度 | 650210 | 音频分析仪校准规范 JJF 1395 | 0.01%～30% | Urel=1.4% |  | 扩项 |
|  |  | 电压 |  |  | 1mV～300V | Urel=0.3% |  |  |
|  |  | 频率 |  |  | 10Hz～200kHz | Urel=0.5% |  |  |
| 190 | 频率表 | 频率 | 660105 | 频率表检定规程 JJG 603 | 45Hz～1kHz | Urel=0.1% |  | 扩项 |
| 191 | 同轴式衰减器 | 衰减电平 | 650117 | 同轴式衰减器检定规程 JJG 387 | (0～100)dB(10Hz-4GHz) | U=(0.3～1.5)dB |  | 扩项 |
| 192 | 电导率仪 | 电导率 | 680709 | 电导率仪检定规程 JJG 376 | (0.05～20000)μS/cm | U=1.2%FS |  | 扩项 |
| 193 | pH计检定仪 | 电压 | 680705 | pH计检定仪检定规程 JJG 919 | (-1000～1000)mV | Urel=0.027% |  | 扩项 |
|  |  | 电阻 |  |  | (0.1～10)GΩ | Urel=0.3% |  |  |
| 194 | 酸度计(pH计) | pH | 680701 | 实验室pH（酸度）计检定规程 JJG 119 | (0～14)pH | 电计部分:U=0.002pH |  | 扩项 |
| 195 | 离子计 | pX | 680703 | 离子计检定规程 JJG757 | (0-14)pX | U=0.002 pX |  | 扩项 |
| 196 | 液相色谱仪 | 最小检测浓度 | 680203 | 液相色谱仪检定规程 JJG 705 | 紫外-可见光、二极管阵列检测器：≤5×10-8g/mL | Urel=6% |  | 扩项 |
|  |  |  |  |  | 荧光检测器：≤5×10-9g/mL | Urel=6% |  |  |
|  |  |  |  |  | 示差折光率、蒸发光散射检测器：≤5×10-6g/mL | Urel=7% |  |  |
| 197 | 原子吸收分光光度计 | 检出限 | 680108 | 原子吸收分光光度计 JJG 694 | 火焰法测铜：≤2×10-2μg/mL | U＝2.5×10-3µg/ mL |  | 扩项 |
|  |  |  |  |  | 石墨炉法测镉：≤4pg | U＝0.2pg |  |  |
| 198 | 电感耦合等离子发射光谱仪 | 检出限 | 680116 | 发射光谱仪检定规程 JJG 768 | Zn：≤3×10-3mg/L | U=2.8×10-4mg/L |  | 扩项 |
|  |  |  |  |  | Ni：≤1×10-2mg/L | U=1.2×10-3mg/L |  |  |
|  |  |  |  |  | Mn:≤2×10-3mg/L | U=2.6×10-4mg/L |  |  |
|  |  |  |  |  | Cr、Cu：≤7×10-3mg/L | U=1×10-3mg/L |  |  |
|  |  |  |  |  | Ba：≤1×10-3mg/L | U=1.6×10-4mg/L |  |  |
| 199 | 气相色谱仪 | 灵敏度 | 680201 | 气相色谱仪检定规程 JJG700 | TCD：≥800 mv·mL/mg | Urel=5% |  | 扩项 |
|  |  | 检出限 |  |  | FID：≤1×10-10g/s | Urel=5% |  |  |
|  |  | 检出限 |  |  | FPD：硫≤5×10-10g/s；磷≤1×10-10g/s | Urel=6% |  |  |
|  |  |  |  |  | NPD：氮≤5×10-12g/s；磷≤1×10-11g/s | Urel=5% |  |  |
|  |  |  |  |  | ECD：≤5×10-12g/mL | Urel=5% |  |  |
| 200 | 波长色散X射线荧光光谱仪 | 计数率 | 680112 | 波长色散X射线荧光光谱仪检定规程 JJG810 | (0～40)% | Urel=10% |  | 扩项 |
| 201 | 能量色散X射线荧光光谱仪 | 检出限 | 680112 | 能量色散X射线荧光光谱仪校准规范 JJF(闽)1047 | (0～1200)mg/kg | Cd:Urel=8% |  | 扩项 |
| 202 | 离子色谱仪 | 最小检出浓度 | 680204 | 离子色谱仪检定规程 JJG 823 | ≤0.02μg/mL | Urel=4.0% |  | 扩项 |
| 203 | 傅里叶变换红外光谱仪 | 波数 | 680107 | 傅立叶变换红外光谱仪校准规范 JJF 1319 | (400～4000)cm-1 | U=0.5cm-1 |  | 扩项 |
| 204 | 原子荧光光度计 | 检出限 | 680110 | 原子荧光光度计 JJG 939-2009 | Sb、As：≤0.4ng | Urel=30% |  | 扩项 |
| 205 | 卡尔费休水分仪 | 水分 | 680507 | 卡尔.费休库仑法微量水分测定仪检定规程 JJG 1044 | 0.1μg～100mg | Urel=1.5% |  | 扩项 |
| 206 | 光源箱 | 光照度 | 670106 | 标准光源箱校准规范 JJF(纺织）055 | (1～10000)lx | Urel=1.7% |  | 扩项 |
|  |  | 分布色温 |  |  | (1500～10000)K | U=20K |  |  |
| 207 | 色温表 | 色温 | 670199 | 色温表检定规程 JJG212 | (2800～7000)K | U=20K |  | 扩项 |
| 208 | 光照度计 | 光照度 | 670105 | 光照度计检定规程 JJG 245 | (5～3000)lx | Urel=1.5% |  | 扩项 |
| 209 | 亮度计 | 亮度 | 670107 | 亮度计检定规程 JJG 211 | (10～1000)cd/m2 | Urel=2.5% |  | 扩项 |
|  |  | 色品坐标 |  |  | x:(0～1)y:(0～1) | U(x)=0.02 U(y)=0.02 |  |  |
| 210 | LED 光电综合测试仪 | 色温 | 670199 | 瞬态光谱仪校准规范 JJF1329,总光通量标准白炽灯检定规程 JJG247 | (1500～6500)K | U=18K |  | 扩项 |
|  |  | 光通量 | 670110 |  | (50～1400) lm | Urel=2.5% |  |  |
| 211 | 澄明度检测仪 | 照度 | 670199 | 澄明度检测仪校准规范 JJF 1287 | (0～5000)lx； | Urel=4.2% |  | 扩项 |
| 212 | 漫透射视觉密度计 | 密度 | 670526 | 漫透射视觉密度计检定规程 JJG920 | D:(0.01～＜2.0) | Urel=2.0% |  | 扩项 |
|  |  |  |  |  | D:(2.0～＜4.0) | Urel=3.0% |  |  |
| 213 | *熔体流动速率仪 | 温度 | 680609 | 熔体流动速率仪检定规程 JJG 878 | (0～400)℃ | U=0.6℃ |  | 扩项 |
|  |  | 流速 |  |  | (0.1～25)g/10min | U=(0.5～0.8)g/10min |  |  |
| 214 | 电线低温冲击试验装置 | 长度 | 701099 | 橡皮塑料电线电缆试验仪器设备检定方法 第2部分：低温冲击试验装置 JB/T4278.2 | (10～200)mm | U=0.10mm |  | 扩项 |
|  |  | 压力 |  |  | (0.1～20)N | U=15mN |  |  |
| 215 | 电线动态曲挠试验机 | 长度 | 620719 | 橡皮塑料电线电缆试验仪器设备检定方法 第3部分：曲挠试验装置 JB/T4278.3 | (10～200)mm | U=0.10mm |  | 扩项 |
|  |  |  |  |  | (200～1000)mm | U=1mm |  |  |
|  |  | 速度 |  |  | (0.2～0.5)m/s | U=0.015m/s |  |  |
|  |  | 压力 |  |  | (0.1～20)N | U=15mN |  |  |
|  |  | 角度 |  |  | （0~45）° | U=0.5° |  |  |
| 216 | 电线耐磨试验机 | 长度 | 620718 | 橡皮塑料电线电缆试验仪器设备检定方法 第4部分：耐磨试验装置 JB/T4278.4 | (10～200)mm | U=0.10mm |  | 扩项 |
|  |  |  |  |  | (200～1000)mm | U=1mm |  |  |
| 217 | 单根绝缘电线电缆垂直燃烧试验装置 | 长度 | 711024 | 橡皮塑料电线电缆试验仪器设备检定方法 第5部分：单根绝缘电线电缆垂直燃烧试验装置 JB/T4278.5 | (1～150)mm | U=0.03mm |  | 扩项 |
|  |  |  |  |  | (200～2000)mm | U=1mm |  |  |
|  |  | 角度 |  |  | (0~45)° | U=0.5° |  |  |
|  |  | 时间 |  |  | (10～60)s | U=1.0s |  |  |
| 218 | 电线低温卷绕试验机 | 长度 | 701099 | 橡皮塑料电线电缆试验仪器设备检定方法 第11部分：低温卷绕试验机 JB/T4278.11 | (1～150)mm | U=0.03mm |  | 扩项 |
|  |  | 转速 |  |  | (10～20)r/min | U=0.15r/min |  |  |
| 219 | 电线高温压力试验装置 | 长度 | 701099 | 橡皮塑料电线电缆试验仪器设备检定方法 第12部分：高温压力试验装置 JB/T4278.12 | (0.50～1.0)mm | U=0.003mm |  | 扩项 |
|  |  | 力值 |  |  | (0.1～20)N | U=15mN |  |  |
| 220 | 导体电阻测量夹具 | 长度 | 701099 | 漆包绕组线试验仪器设备检定方法 第2部分：直流电阻试验装置 JB/T4279.2 | (1～150)mm | U=0.03mm |  | 扩项 |
|  |  |  |  |  | (200～2000)mm | U=0.2mm |  |  |
| 221 | 伸长试验仪 | 长度 | 600299 | 漆包绕组线试验仪器设备检定方法 第3部分： 伸长试验仪 JB/T4279.3 | (10～300)mm | U=0.10mm |  | 扩项 |
|  |  | 伸长率 |  |  | (5～50)% | U=0.2% |  |  |
|  |  | 速度 |  |  | (100～500)mm/min | U=10mm/min |  |  |
| 222 | 回弹角试验仪 | 长度 | 701099 | 漆包绕组线试验仪器设备检定方法 第4部分：回弹角试验仪 JB/T4279.4 | (5～50)mm | U=0.03mm |  | 扩项 |
|  |  | 速度 |  |  | (5～10)r/min | U=0.15r/min |  |  |
|  |  | 压力 |  |  | (0.10～2)N | U=0.005N |  |  |
|  |  |  |  |  | (2～20)N | U=0.05N |  |  |
| 223 | 剥离试验仪 | 长度 | 701099 | 漆包绕组线试验仪器设备检定方法 第6部分：剥离试验仪 JB/T4279.6 | (495～505)mm | U=1mm |  | 扩项 |
|  |  | 转速 |  |  | (50～100)r/min | U=2r/min |  |  |
|  |  | 压力 |  |  | (25～100)N | Urel=3% |  |  |
| 224 | *软化击穿试验仪 | 交流电压 | 701099 | 漆包绕组线试验仪器设备检定方法 第8部分:软化击穿试验仪 JB/T4279.8 | (20～200)V (45Hz～65Hz) | Urel=1.0% |  | 扩项 |
|  |  | 交流电流 |  |  | (2～20)mA (45Hz～65Hz) | Urel=1.5% |  |  |
|  |  | 温度 |  |  | (50～500)℃ | U=1.0℃ |  |  |
|  |  | 时间间隔 |  |  | (10～3600)s | U=0.5s |  |  |
|  |  | 负荷重力 |  |  | (1～5)N | U=0.015N |  |  |
|  |  |  |  |  | (5～50)N | U=0.05N |  |  |
| 225 | *单向刮漆试验仪 | 长度 | 701099 | 漆包绕组线试验仪器设备检定方法 第9部分:单向刮漆试验仪 JB/T4279.9 | (1～25)mm | U=0.05mm |  | 扩项 |
|  |  | 移动速度 |  |  | (50～500)mm/min | U=10mm/min |  |  |
|  |  | 直流电压 |  |  | (2～20)V | U=0.10V |  |  |
|  |  | 直流电流 |  |  | (5～50)mA | U=0.15mA |  |  |
|  |  | 力值 |  |  | (0.2～20)N | U=0.02N |  |  |
| 226 | *耐溶剂试验仪 | 温度 |  | 漆包绕组线试验仪器设备检定方法 第10部分：耐溶剂试验仪 JB/T4279.10 | (60～100)℃ | U=0.5℃ |  | 扩项 |
|  |  | 角度 |  |  | (10～90)° | U=1.0° |  |  |
|  |  | 力值 |  |  | (1～10)N | U=0.05N |  |  |
| 227 | *击穿电压试验仪 | 交流电压 | 701099 | 漆包绕组线试验仪器设备检定方法 第11 部分 击穿电压试验仪 JB/T4279.11 | (0.1～10)kV | Urel=1.5% |  | 扩项 |
|  |  | 交流电流 |  |  | (2～10)mA | Urel=1.5% |  |  |
|  |  | 升压速度 |  |  | (10～500)V/s | Urel=3% |  |  |
| 228 | *低压漆膜连续性试验仪 | 长度 | 701099 | 漆包绕组线试验仪器设备检定方法 第12部分：低压漆膜连续性试验仪 JB/T4279.12 | (1～150)mm | U=0.03mm |  | 扩项 |
|  |  | 速度 |  |  | (10～20)m/min | U=0.5m/min |  |  |
|  |  | 直流电压 |  |  | (10～100)V | Urel=0.5% |  |  |
| 229 | *焊锡试验仪 | 温度 |  | 漆包绕组线试验仪器设备检定方法 第14部分：焊锡试验仪 JB/T4279.14 | (100～500)℃ | U=(1.0～1.5)℃ |  | 扩项 |
| 230 | *砂浆稠度仪 | 长度 | 700810 | 非金属建材塑限测定仪校准规范 JJF 1090 | (0～150)mm | U=0.05mm |  | 扩项 |
|  |  | 角度 |  |  | (0～45)º | U=0.06° |  |  |
|  |  | 质量 |  |  | (0～300)g | U=0.3g |  |  |
| 231 | *水泥净浆标准稠度与凝结时间测定仪 | 角度 | 700809 | 非金属建材塑限测定仪校准规范 JJF1090 | (0～45)º | U=0.06° |  | 扩项 |
|  |  | 长度 |  |  | Φ(0～25)mm | U=0.02mm |  |  |
|  |  |  |  |  | (0～100)mm | U=0.05mm |  |  |
|  |  | 质量 |  |  | (0～300)g | U=0.3g |  |  |
| 232 | *离心机 | 转速 | 621112 | 离心式恒加速度试验机检定规程 JJG 972 | （1～20000）r/min | Urel=1.0% |  |  |
| 233 | 条式水平仪 | 角度 | 600313 | 框式、条式水平仪校准规范 JJF1084 | （0.02～0.08）mm/m | Urel=5.8% |  |  |
| 234 | 框式水平仪 | 角度 | 600313 | 框式、条式水平仪校准规范 JJF1084 | （0.02～0.08）mm/m | Urel=5.8% |  | 扩项 |
| 235 | 金相显微镜 | 长度 | 600699 | 金相显微镜检定规程 JJG(教委)012 | 4x～100x | Urel=3% |  | 扩项 |